WALK - IN – CHAMBER

WALK - IN - STABILITY CHAMBER :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8418 |
| Control Condition | Long term testing - + 250C : RH 60% OR + 300C : RH 75% |
|  | Accelerated testing - + 400C : RH 75% |
|  | Intermediate testing - + 300C : RH 65% |
| Construction | - Double walled insulated PUF modular panels for easy assembling at site. - Leak proof door with full gasket & air tight mechanical lock ensures no leakage through chamber. - Observation window (30 x 45 cm) for viewing sample inside the chamber. - The unit is provided with Perforated Stainless Steel Trays. |
| Insulation | 80 mm thick best quality of PUF insulation. |
| Air Circulation | Forced air to assure maximum uniformity of Temperature inside the chamber. |
| Heating System | High quality S. S. Tubular Heaters (made of Nichrome wire sheathed in S.S. tube). Special arrangement of Heaters avoids direct radiation on stored samples. |
| Humidity System | Non-condensing type steam injection system with water level switch. Humidity System provided with a low pressure Boiler tank & Reservoir tank dully Argon welded with Float valve. Low water level cut off switch ensures sufficient water level in Boiler tank. DM Water is preferred. In Built Stand by Humidity system with auto changes over facility. |
| Refrigeration System | By means of CFC Free hermitically sealed Emerson Copland make compressor for better cooling with ECO friendly refrigerant. The cooling system is made up of split type so as one can place the compressor coupled with condenser unit in well – ventilated area and avoid hot air throw in the Stability room. In built Stand by Refrigeration system with auto change over facility. |
| Safety features | - Audio Visual alarms are generated for temperature /humidity variation. - Float switch to cutoff the wet heater supply in case of low water level in the boiler tank with audio & visual alarm. - Control system should take care for avoiding overshooting or undershooting of control parameter in worst case it will give audio visual alarm. - Safety thermostat will cut off the main contactor through PLC in case of variation in the set temperature conditions. - Compressor time delay for compressor safety. - Emergency stop button inside the chamber. |
| Features | - Negligible Maintenance – The sturdy design and long – lasting finish ensures minimum maintenance for years. - Minimal Heat Transfer – The tongue and groove edges of all panels are molded to close tolerance ensuring tight fit between panels. This prevents condensation from forming on the foam and minimize heat transfer. - Quality Accessories – Hardware like hinges, handles etc. on panels from best indigenous / foreign suppliers. |
| Illumination | Interior illumination for Working area. |
| Control System | - Microprocessor based PLC (up to 5000 readings) with 7” Touch screen colour HMI, which gives the direct display of process value & set value of temperature & Humidity in the chamber. - High speed Ethernet based communication. |
| Sensor | Temperature & humidity sensor (Transmitter type 4-20mA) will accurately measure the temperature & Humidity. |
| Scanner | 4 Point Temperature & 4 Point Humidity Scanner complete with Sensors. |
| Supply | 230 V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304 |
| Temperature Range | 200 C to 600 C |
| Temperature Accuracy | ± 0.20 C |
| Temperature Uniformity | + 2 0C |
| Humidity Range | 40% to 95% Rh |
| Humidity Accuracy | + 2% Rh |
| Humidity Uniformity | + 3% Rh |

| Models | Cap. Liters | Internal size H x W x D Cms. | External Size H x W x D Cms. | Trays Nos. | Tray Size Cm. | Inner & Outer SS 304 (Amt. in Rs.) |
| --- | --- | --- | --- | --- | --- | --- |
| TI-710 WIC | 5400 | 200 x 136 x 200 | 216 x 152 x 216 | 24 | 87 x 30 | PRICE ON REQUEST |
| TI-711 WIC | 6400 | 240 x 136 x 200 | 256 x 152 x 216 | 32 | 87 x 30 | PRICE ON REQUEST |
| TI-712 WIC | 8000 | 200 x 200 x 200 | 216 x 216 x 216 | 24 | 87 x 60 | PRICE ON REQUEST |
| TI-713 WIC | 9600 | 240 x 200 x 200 | 256 x 216 x 216 | 32 | 87 x 60 | PRICE ON REQUEST |
| TI-714 WIC | 15000 | 240 x 250 x 250 | 256 x 266 x 266 | 48 | 72 x 80 | PRICE ON REQUEST |
| TI-715 WIC | 18000 | 240 x 250 x 300 | 256 x 266 x 316 | 48 | 90 x 80 | PRICE ON REQUEST |

ERECTION CHARGES WILL BE COST EXTRA RS. 50,000/-

Documentation :  We will provide IQ, OQ & PQ Documents in soft copy along with Material Test report &

Calibration certificate .

Validation         :  24 hrs. Loaded condition OR unloaded condition on one cycle - INCLUSIVE.

(If any additional Validation set point will be charged Extra)

Cooling Piping  :  For Walk in Chamber, the distance between the chamber & the compressor unit should be

max. 10 mtrs.  Further extra length will be chargeable @ Rs.1000/- per mtr. Inclusive of

piping & insulation.

OPTIONALS :-

| 1 | CFR 21 Part 11 compliance window based communication software for data management. Complete with Ethernet multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 30,000 |
| --- | --- | --- |
| 2 | Mobile alert system (Only for 5 nos. Mobile device) | 12,000 |
| 3 | Security System (Door Access) | 16,500 |
| 4 | Scanner : 8 point temperature & 8 point Humidity complete with sensors. | 1,25,000 |

Requirement for Installation :

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally levelled.

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not use pump water, pressured water or hot water for water connection. Water supply from an overhead

- water tank where the water pressure should not be more than 15 PSI is recommended.

- Use of Soft water or D. M. water for humidity formation is recommended. Use of hard or salty water will

damage the boiler heater.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

WALK - IN - BOD INCUBATOR :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8418 |
| Construction | - Double walled insulated PUF modular panels for easy assembling at site. - Leak proof door with full gasket & air tight mechanical lock ensures no leakage through chamber. - Observation window (30 x 45 cm) for viewing sample inside the chamber. - The unit is provided with Perforated Stainless Steel Trays. |
| Insulation | 80 mm thick best quality of PUF insulation. |
| Air Circulation | Forced air to assure maximum uniformity of Temperature inside the chamber. |
| Heating System | High quality S. S. Tubular Heaters (made of Nichrome wire sheathed in S.S. tube). Special arrangement of Heaters avoids direct radiation on stored samples. |
| Refrigeration System | By means of CFC Free hermitically sealed Emerson Copland make compressor for better cooling with ECO friendly refrigerant. The cooling system is made up of split type so as one can place the compressor coupled with condenser unit in well – ventilated area and avoid hot air throw in the Stability room. In built Stand by Refrigeration system with auto change over facility. |
| Safety features | - Audio Visual alarms are generated for temperature variation. - Control system should take care for avoiding overshooting or undershooting of control parameter in worst case it will give audio visual alarm. - Safety thermostat will cut off the main contactor through PLC in case of variation in the set temperature conditions. - Compressor time delay for compressor safety. - Emergency stop button inside the chamber. |
| Features | - Negligible Maintenance – The sturdy design and long – lasting finish ensures minimum maintenance for years. - Minimal Heat Transfer – The tongue and groove edges of all panels are molded to close tolerance ensuring tight fit between panels. This prevents condensation from forming on the foam and minimize heat transfer. - Quality Accessories – Hardware like hinges, handles etc. on panels from best indigenous / foreign suppliers. |
| Illumination | Interior illumination for Working area. |
| Control System | - Microprocessor based PLC (up to 5000 readings) with 7” Touch screen colour HMI, which gives the direct display of process value & set value of temperature in the chamber. - High speed Ethernet based communication. |
| Sensor | Temperature sensor (PT 100) will accurately measure the temperature. |
| Scanner | 4 Point Temperature Scanner complete with Sensors. |
| Supply | 230 V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304 |
| Temperature Range | 100 C to 500 C |
| Temperature Accuracy | ± 0.50 C |
| Temperature Uniformity | + 2 0C |

| Models | Cap. Ltr. | Internal size H x W x D Cm. | External Size H x W x D Cm. | Tray Nos. | Trays Size | Inner & Outer SS 304 (Amt. in Rs.) |
| --- | --- | --- | --- | --- | --- | --- |
| TI-500 WIC | 5400 | 200 x 136 x 200 | 216 x 152 x 216 | 24 nos. | 87 x 30 | PRICE ON REQUEST |
| TI-501 WIC | 6400 | 240 x 136 x 200 | 256 x 152 x 216 | 32 nos. | 87 x 30 | PRICE ON REQUEST |
| TI-502 WIC | 8000 | 200 x 200 x 200 | 216 x 216 x 216 | 24 nos. | 87 x 60 | PRICE ON REQUEST |
| TI-503 WIC | 9600 | 240 x 200 x 200 | 256 x 216 x 216 | 32 nos. | 87 x 60 | PRICE ON REQUEST |
| TI-504 WIC | 15000 | 240 x 250 x 250 | 256 x 266 x 266 | 48 nos. | 72 x 80 | PRICE ON REQUEST |
| TI-505 WIC | 18000 | 240 x 250 x 300 | 256 x 266 x 316 | 48 nos. | 90 x 80 | PRICE ON REQUEST |

ERECTION CHARGES WILL BE COST EXTRA RS. 50,000/-

Documentation  :  We will provide IQ, OQ & PQ Documents in soft copy along with Material Test report &

Calibration certificate .

Validation         :  24 hrs. Loaded condition OR unloaded condition on one cycle (INCLUSIVE).

(If any additional Validation set point will be charged Extra.)

Cooling Piping   :  For Walk in Chamber, the distance between the chamber & the compressor unit should be

max. 10 mtrs.  Further extra length will be chargeable @ Rs.1000/- per mtr. Inclusive of

piping & insulation.

OPTIONALS :-

| 1 | CFR 21 Part 11 compliance window based communication software for data management. Complete with Ethernet multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 30,000 |
| --- | --- | --- |
| 2 | Mobile alert system (Only for 5 nos. Mobile device). | 12,000 |
| 3 | Security System (Door Access) | 16,500 |
| 4 | Scanner : 8 point temperature complete with sensors. | 33,000 |

Requirement for Installation :

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally levelled.

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not use pump water, pressured water or hot water for water connection. Water supply from an overhead

water tank where the water pressure should not be more than 15 PSI is recommended.

- Use of Soft water or D. M. water for humidity formation is recommended. Use of hard or salty water will

damage the boiler heater.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

WALK - IN - LAB REFRIGERATOR :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8418 |
| Construction | - Double walled insulated PUF modular panels for easy assembling at site. - Leak proof door with full gasket & air tight mechanical lock ensures no leakage through chamber. - Observation window (30 x 45 cm) for viewing sample inside the chamber. - The unit is provided with Perforated Stainless Steel Trays. |
| Insulation | 80 mm thick best quality of PUF insulation. |
| Air Circulation | Forced air to assure maximum uniformity of Temperature inside the chamber. |
| Heating System | High quality S. S. Tubular Heaters (made of Nichrome wire sheathed in S.S. tube). Special arrangement of Heaters avoids direct radiation on stored samples. |
| Refrigeration System | By means of CFC Free hermitically sealed Emerson Copland make compressor for better cooling with ECO friendly refrigerant. The cooling system is made up of split type so as one can place the compressor coupled with condenser unit in well – ventilated area and avoid hot air throw in the Stability room In built Stand by Refrigeration system with auto change over facility. |
| Safety features | - Audio Visual alarms are generated for temperature variation. - Control system should take care for avoiding overshooting or undershooting of control parameter in worst case it will give audio visual alarm. - Safety thermostat will cut off the main contactor through PLC in case of variation in the set temperature conditions. - Compressor time delay for compressor safety. - Emergency stop button inside the chamber. |
| Features | - Negligible Maintenance – The sturdy design and long – lasting finish ensures minimum maintenance for years. - Minimal Heat Transfer – The tongue and groove edges of all panels are molded to close tolerance ensuring tight fit between panels. This prevents condensation from forming on the foam and minimize heat transfer. - Quality Accessories – Hardware like hinges, handles etc. on panels from best indigenous / foreign suppliers. |
| Illumination | Interior illumination for Working area. |
| Control System | - Microprocessor based PLC (up to 5000 readings) with 7” Touch screen colour HMI, which gives the direct display of process value & set value of temperature in the chamber. - High speed Ethernet based communication. |
| Sensor | Temperature sensor (PT 100) will accurately measure the temperature. |
| Scanner | 4 Point Temperature Scanner complete with Sensors. |
| Supply | 230 V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304. |
| Temperature Range | 20 C to 80 C |
| Temperature Accuracy | ± 0.50 C |
| Temperature Uniformity | + 2 0C |

| Models | Cap. Ltrs. | Internal size H x W x D in Cm. | External Size H x W x D in Cm. | Tray Nos. | Trays Size | Inner & outer S S 304 (Amt. in Rs) |
| --- | --- | --- | --- | --- | --- | --- |
| TI-810 WIC | 5400 | 200 x 136 x 200 | 216 x 152 x 216 | 24 nos. | 87 x 30 | PRICE ON REQUEST |
| TI-811 WIC | 6400 | 240 x 136 x 200 | 256 x 152 x 216 | 32 nos. | 87 x 30 | PRICE ON REQUEST |
| TI-812 WIC | 8000 | 200 x 200 x 200 | 216 x 216 x 216 | 24 nos. | 87 x 60 | PRICE ON REQUEST |
| TI-813 WIC | 9600 | 240 x 200 x 200 | 256 x 216 x 216 | 32 nos. | 87 x 60 | PRICE ON REQUEST |
| TI-814 WIC | 15000 | 240 x 250 x 250 | 256 x 266 x 266 | 48 nos. | 72 x 80 | PRICE ON REQUEST |
| TI-815 WIC | 18000 | 240 x 250 x 300 | 256 x 266 x 316 | 48 nos. | 90 x 80 | PRICE ON REQUEST |

ERECTION CHARGES WILL BE COST EXTRA RS. 50,000/-

Documentation  :  We will provide IQ, OQ & PQ Documents in soft copy along with Material Test report &

Calibration certificate .

Validation           :  24 hrs. Loaded condition OR unloaded condition on one cycle. (INCLUSIVE)

(If any additional Validation set point will be charged Extra.

Cooling Piping    :  For Walk in Chamber, the distance between the chamber & the compressor unit should be

max. 10 mtrs.  Further extra length will be chargeable @ Rs.1,000/- per mtr. Inclusive of

piping & insulation.

OPTIONALS :-

| 1 | CFR 21 Part 11 compliance window based communication software for data management. Complete with Ethernet multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 30,000 |
| --- | --- | --- |
| 2 | Mobile alert system (Only for 5 nos. Mobile device). | 12,000 |
| 3 | Security System (Door Access) | 16,500 |
| 4 | Scanner : 8 point temperature complete with sensors. | 33,000 |

Requirement for Installation :

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally levelled.

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not use pump water, pressured water or hot water for water connection. Water supply from an overhead

water tank where the water pressure should not be more than 15 PSI is recommended.

- Use of Soft water or D. M. water for humidity formation is recommended. Use of hard or salty water will

damage the boiler heater.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

PRODUCT RANGE (COOLING)

STABILITY CHAMBER :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Control condition | As per ICH guidelines to control following conditions: At 25 0C : RH = 60% At 30 0C : RH = 65% At 30 0C : RH = 75% At 40 0C : RH = 75% |
| Construction | Double walled construction with backside triple. |
| Insulation | 3” Thick PUF Insulation ensures stable temperature with humidity & reduced energy consumption. |
| Air Circulation | Motorized Blower at back side of the chamber develops unique air flow system which ensures maximum uniform temperature & humidity condition inside the chamber. Unique air flow assures quick recovery after door openings. |
| Blower motor | 1/12 HP, TEFC , F-class insulation, Single phase, 1440 RPM, 230 Volts |
| Heating System | High quality U shape S.S. Tubular Heaters are used for better heating. |
| Humidity System | Steam injection system (Non condensing type) & reservoir tank with water level arrangement with indicator. |
| Safety features | - In case of Temperature variation, Temperature deviation alarm will active & give audio visual alarm. - In case of Humidity variation, Humidity deviation alarm will active & give audio visual alarm. - Cut off the supply in case temperature goes above 600 C & activate safety alarm. - Float switch to cut off the wet heater supply in case of low water level in the Boiler tank and give Audio & Visual alarm. - Compressor ON delay timer (2 min.) to safe guard the compressor. - In built over load protector provided for hermetically sealed compressor. |
| Features | - Fiber/PU Wheel with front lockable for easy movement. - MCB for electrical safety - Adjustable tray height arrangements. |
| Refrigeration System | By means of CFC Free hermetically sealed Emerson Copland make compressor with R 134 A for better cooling with ECO friendly refrigerant with time delay to safe guard compressor system. |
| Validation Port | 50 mm validation port hole with Silicon rubber seal to insert sensors for validation. |
| Observation Door | Full view observation Acrylic Door / Glass Door with gasket to observe sample inside the chamber. |
| Lighting | Interior illumination for Working area. |
| Trays | Specially designed stainless steel rod trays ensure uniform temperature distribution. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model. | Inside S.S. 304 & Outer Galvanized Iron (GI) sheet with powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304 |
| G.M.P. Model | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range | 100 C to 600 C |
| Temperature Accuracy | ± 0.20 C |
| Temperature Uniformity | ± 1 0 C |
| Humidity Range | 40% to 95% |
| Humidity Accuracy | + 2 % |
| Humidity Uniformity | + 3 % |
| Temperature Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value. |
| Humidity Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value with direct % RH. Imported RH Sensor - Box type Rotronic Sensor with 95% RH |

| Models | Capacity Liters | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | G.M.P. Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TI-710 | 100 | 50 x 50 x 40 | 2 | 1,50,700 | 1,74,900 | 1,89,800 | 5,000 |
| TI-711 | 170 | 70 x 60 x 40 | 2 | 1,62,800 | 2,01,500 | 2,20,800 | 5,000 |
| TI-712 | 200 | 60 x 60 x 60 | 2 | 1,64,800 | 2,03,700 | 2,23,300 | 5,500 |
| TI-713 | 285 | 80 x 60 x 60 | 3 | 1,75,700 | 2,23,300 | 2,44,600 | 5,500 |
| TI-714 | 325 | 90 x 60 x 60 | 4 | 1,90,100 | 2,44,600 | 2,68,200 | 6,000 |
| TI-715 | 400 | 110 x 60 x 60 | 4 | 2,04,600 | 2,67,800 | 2,93,600 | 8,000 |
| TI-716 | 600 | 125 x 70 x 70 | 4 | 2,52,700 | 2,95,600 | 3,22,100 | 10,000 |
| TI-717 | 800 | 125 x 80 x 80 | 5 | 2,82,700 | 3,28,200 | 3,67,800 | 10,800 |
| TI-718 | 1000 | 157 x 80 x 80 | 5 | 3,02,700 | 3,74,500 | 4,10,900 | 12,000 |
| TI-719 | 1500 Double Door | 146 x 150 x 69 | 10 | PRICE ON REQUEST |  |  | 14,000 |
| TI-720 | 2000 Double Door | 146 x 150 x 92 | 10 | PRICE ON REQUEST |  |  | 15,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

ERECTION CHARGES WILL BE EXTRA RS.50,000/- FOR (TI – 719 & 720 models are BIG & assembled type on SITE)

Note	               :  Two blowers for 400 ltrs to 2000 ltrs. Capacity chambers.

Validation & Documentation  :  2 hrs. Loaded OR unloaded condition on one cycle. (4 or 8 temp. & humidity

probes). Along with IQ, OQ & PQ Documents.

Cost of Validation	               :  EXTRA Rs.  15,000/- + GST @ 18% Extra.

OPTIONALS :-

| A] | Complete PLC system (with touch screen HMI) :- |  |
| --- | --- | --- |
|  | PLC for auto changeover for standby refrigeration & humidity system with 7” Touch screen color HMI. | 60,000 |
|  | 21 CFR Part 11 compliance Ethernet based communication software for data management with multi dropping monitor 8 chambers to one software. | 30,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 100 ltr. to 325 ltr. | 40,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 400 ltr. to 600 ltr. | 50,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 800 ltr. to 2000 ltr. | 65,000 |
|  | Ready to use stand by humidity system in case of regular system fails. | 8,000 |
|  | Scanner: 4 point temperature & 4 point Humidity complete with Rotronic sensors. | 88,000 |
|  | Scanner: 8 point temperature & 8 point Humidity complete with Rotronic sensors. | 1,95,000 |
|  | Mobile alert system (Only for 5 nos. Mobile device) To be ordered with PLC System | 12,000 |
|  | Security System (Door Access) To be ordered with PLC System | 16,500 |
| B] | DATA LOGGER SYSTEM :- |  |
|  | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. | 22,000 |
|  | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. | 20,000 |
| C] | SCANNER |  |
| 1 | Scanner: 4 point temperature & 4 point Humidity complete with Rotronic sensors. | 1,13,000 |
| 2 | Scanner: 8 point temperature & 8 point Humidity complete with Rotronic sensors. | 1,75,000 |
| D] | OTHERS |  |
| 1 | Imported RH Sensor European Make, Rotronic Switzerland for higher accuracy & RH 95% | 18,000 |
| 2 | Safety digital temperature controller with accessories | 6,000 |
| 3 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 6 | Stainless steel rod trays : 100 ltr. to 400 ltr. | 3,000 |
| 7 | Stainless steel rod trays : 600 ltr. to 2000 ltrs. | 4,500 |
| 8 | 30 Ltr S S Tank to store water with Lid & 2 nos. taps connection along with MS Stand for Tank | 12,000 |

Requirements for Installation :

      - PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

      - Place the chamber where ventilation is good.  The chamber must be operated in an ambient of 220C ~ 300C.

      - Leave minimum 1ft.space from the sides and leave 2 ft. space from the backside.

      - Place the chamber horizontally leveled.

      - Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a voltage stabilizer (KVA).

      - Proper grounding of the chamber to get satisfactory operation of chamber.

      - Do not use pump water, pressured water or hot water for water connection. Water supply from an overhead water tank

where the water pressure should not be more than 15 PSI is recommended.

- Use of Soft water or D. M. water for humidity formation is recommended. Use of hard or salty water will damage the boiler heater.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

PHOTOSTABILITY CHAMBER :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Photo stability Effects | Fluorescent tube for day light effect and UV tubes for Ultra Violet ray in accordance to the ICH guidelines requirement of a lightning energy mounting to 1.2 mil lux per hours and integrated UV energy share of 200 Watt Hours per mtr sqr. |
| Construction | Double walled construction with backside triple. |
| Insulation | 3”Thick PUF Insulation ensures stable temperature with humidity & reduced energy consumption. |
| Air Circulation | Motorized Blower at back side of the chamber develops unique air flow system which ensures maximum uniform temperature & humidity condition inside the chamber. Unique air flow assures quick recovery after door openings. |
| Blower motor | 1/12 HP, TEFC , F-class insulation, Single phase, 1440 RPM, 230 Volts |
| Heating System | High quality S.S. Tubular Heaters are used for better heating. |
| Humidity System | Steam injection humidity system none condensing, with water level arrangement. |
| Safety features | - In case of Temperature variation, Temperature deviation alarm will active & give audio visual alarm. - In case of Humidity variation, Humidity deviation alarm will active & give audio visual alarm. - Cut off the supply in case temperature goes above 600 C & activate safety alarm. - Float switch to cutoff the wet heater supply in case of low water level in the Boiler tank. Audio & Visual alarm. - Compressor ON delay timer (2 min.) to safe guard the compressor. - In built over load protector provided for hermetically sealed compressor. - After opening the Door UV light automatic switch off. |
| Features | - Fiber/PU Wheel with front lockable for easy movement. - MCB for electrical safety - Adjustable tray height arrangements. |
| Refrigeration System | By means of CFC Free hermitically sealed Emerson Copland make compressor with R 134 A for better cooling with ECO friendly refrigerant with time delay to safe guard compressor system. |
| Validation Port | 50 mm validation port hole with Silicon rubber seal to insert sensors for validation. |
| Observation Door | Full view observation Toughened Glass door with gasket to observe sample inside the chamber. |
| Trays | Specially designed stainless steel rod trays ensure uniform temperature distribution. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 & Outer Galvanized Iron (GI) sheet with powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304 |
| G.M.P. Model | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range | 200 C to 500 C |
| Temperature Accuracy | ± 0.20 C |
| Temperature Uniformity | ± 10 C |
| Humidity Range | 40% to 90% |
| Humidity Accuracy | + 2% |
| Humidity Uniformity | ± 3% |
| Temperature Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value. |
| Humidity Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value with direct % RH. Imported RH Sensor Box type Rotronic Sensor with 95% RH. |

| Models | Capacity Liters | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | G.M.P. Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TI-910 | 120 | 50 x 60 x 40 | 1 | 1,56,200 | 1,88,900 | 1,93,100 | 5,000 |
| TI-911 | 170 | 70 x 60 x 40 | 1 | 1,69,600 | 2,17,000 | 2,22,300 | 5,000 |
| TI-912 | 200 | 60 x 60 x 60 | 1 | 1,70,700 | 2,13,600 | 2,19,500 | 5,500 |
| TI-913 | 285 | 80 x 60 x 60 | 2 | 1,96,900 | 2,43,200 | 2,50,500 | 5,500 |
| TI-914 | 325 | 90 x 60 x 60 | 2 | 2,02,000 | 2,54,300 | 2,61,700 | 6,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

Validation & Documentation :  2 hrs. Loaded OR unloaded condition on one cycle. (4 or 8 temp. & humidity

probes). Along with IQ, OQ & PQ Documents.

Cost of Validation	               :  EXTRA Rs.  15,000/- + GST @ 18% Extra.

OPTIONALS :-

| A] | Complete PLC system (with touch screen HMI) :- |  |
| --- | --- | --- |
|  | PLC for auto changeover for standby refrigeration & humidity system with 7” Touch screen color HMI. | 60,000 |
|  | 21 CFR Part 11 compliance Ethernet based communication software for data management with multi dropping monitor 8 chambers to one software. | 30,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 100 ltr. to 325 ltr | 40,000 |
|  | Ready to use stand by humidity system in case of regular system fails. | 8,000 |
|  | Scanner: 4 point temperature & 4 point Humidity complete with Rotronic sensors. | 88,000 |
|  | Scanner: 8 point temperature & 8 point Humidity complete with Rotronic sensors. | 1,95,000 |
|  | Mobile alert system (Only for 5 nos. Mobile device) To be ordered with PLC System | 12,000 |
|  | Security System (Door Access) To be ordered with PLC System | 16,500 |
| B] | DATA LOGGER SYSTEM :- |  |
|  | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. | 22,000 |
|  | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. | 20,000 |
| C] | SCANNER |  |
| 1 | Scanner: 4 point temperature & 4 point Humidity complete with Rotronic sensors. | 1,13,000 |
| 2 | Scanner: 8 point temperature & 8 point Humidity complete with Rotronic sensors. | 1,75,000 |
| D] | OTHERS |  |
| 1 | Digital UV Meter To be ordered with PLC System | 45,000 |
| 2 | Digital Lux Meter To be ordered with PLC System | 13,000 |
| 3 | Imported RH Sensor European Make, Rotronic Switzerland for higher accuracy & RH 95% | 18,000 |
| 4 | Safety digital temperature controller with accessories | 6,000 |
| 5 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 6 | Stainless steel rod trays : 100 ltr. to 400 ltr. | 3,000 |
| 7 | 30 Ltr S S Tank to store water with Lid & 2 nos. taps connection along with MS Stand for Tank | 12,000 |

Requirements for Installation :

- PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

- Place the chamber where ventilation is good.  The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2 ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not use pump water, pressured water or hot water for water connection. Water supply from an overhead

water tank where the water pressure should not be more than 15 PSI is recommended.

- Use of Soft water or D. M. water for humidity formation is recommended. Use of hard or salty water will

damage the boiler heater.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

HUMIDITY OVEN

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Construction | Double walled construction with backside triple. |
| Insulation | 3”Thick PUF Insulation ensures stable temperature with humidity & reduced energy consumption. |
| Air Circulation | Motorized Blower at back side of the chamber develops unique air flow system which ensures maximum uniform temperature & humidity condition inside the chamber. Unique air flow assures quick recovery after door openings. |
| Blower motor | 1/12 HP, TEFC , F-class insulation, Single phase, 1440 RPM, 230 Volts |
| Heating System | High quality U shape S.S. Tubular Heaters are used for better heating. |
| Humidity System | Steam injection humidity system none condensing, with water level arrangement. |
| Safety features | - Additional Safety thermostat to cut of heater supply in case temperature of overshoot. - Float switch to cutoff the wet heater supply in case of low water level in the Boiler tank. Audio & Visual alarm. |
| Features | - Fiber/PU Wheel with front lockable for easy movement. - MCB for electrical safety - Adjustable tray height arrangements. |
| Observation Door | Full view observation Toughened Glass door with gasket to observe sample inside chamber |
| Lighting - not provided | Interior illumination for Working area is not provide as this oven works on highest humidity set points & tube connections get fused. |
| Trays | Specially designed stainless steel rod trays ensure uniform temperature distribution. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 & Outer Galvanized Iron (GI) sheet with powder coated. |
| Temperature Range | 50 C above ambient to 600 C |
| Temperature Accuracy | ± 0.50 C |
| Temperature Uniformity | ± 2 0 C |
| Humidity Range | Above 90% to 95% (Please confirm Set point) |
| Humidity Accuracy | + 3 % |
| Humidity Uniformity | + 5 % |
| Temperature Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value. |
| Humidity Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value with direct % RH. Imported RH Sensor European Make, Rotronic for higher accuracy & RH95% |

| Models | Capacity Liters | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-701 | 100 | 50 x 50 x 40 | 2 | 1,16,800 | 5,000 |
| TI-702 | 170 | 70 x 60 x 40 | 2 | 1,30,200 | 5,000 |
| TI-703 | 200 | 60 x 60 x 60 | 2 | 1,37,400 | 5,500 |
| TI-704 | 285 | 80 x 60 x 60 | 3 | 1,47,700 | 5,500 |
| TI-705 | 325 | 90 x 60 x 60 | 4 | 1,55,600 | 6,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

Requirements for Installation :

- PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

- Place the chamber where ventilation is good.  The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2 ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not use pump water, pressured water or hot water for water connection. Water supply from an overhead

water tank where the water pressure should not be more than 15 PSI is recommended.

- Use of Soft water or D. M. water for humidity formation is recommended. Use of hard or salty water will

damage the boiler heater.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

BOD INCUBATOR :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | Double walled construction with backside triple. |
| Insulation | 3” Thick PUF Insulation ensures stable temperature & reduced energy consumption |
| Air Circulation | Motorized Blower at back side of the chamber develops unique air flow system which maximum uniform temperature condition inside the chamber. Unique air flow assures quick recovery after door openings. |
| Blower motor | 1/12 HP, TEFC , F-class insulation, Single phase, 1440 RPM, 230 Volts |
| Heating System | High quality S.S. Tubular Heaters are used for better heating. |
| Observation Door | Full view observation Acrylic Door / Glass Door with gasket to observe sample inside the chamber. |
| Refrigeration system | By means of CFC Free hermitically sealed Emerson Copland make compressor with R 134 A for better cooling with ECO friendly refrigerant with time delay to safe guard compressor system. |
| Validation Port | 50 mm validation port hole with Silicon rubber seal to insert sensors for validation. |
| Lighting | Interior illumination for working area. |
| Safety features | - In case of Temperature variation, Temperature deviation alarm will active & give audio visual alarm. - Safety Thermostat Cut off the supply in case temperature goes above 600 C & activate safety alarm. - Compressor ON delay timer (2 min.) to safe guard the compressor. - In built over load protector provided for hermetically sealed compressor. |
| Features | - Fiber/PU Wheel with front lockable for easy movement. - MCB for electrical safety. - Adjustable tray height arrangements. |
| Trays | Specially designed stainless steel rod trays ensure uniform temperature distribution. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 & Outer Galvanized Iron (GI) sheet with powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304 |
| G.M.P. Model | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range | 50 C to 600 C |
| Temperature Accuracy | ± 0.50 C |
| Temperature Uniformity | ± 20 C |
| Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value. |

| Models | Cap. Liters | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | G.M.P.Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TI-500 | 100 | 50 x 50 x 40 | 2 | 1,07,500 | 1,36,200 | 1,42,100 | 5,000 |
| TI-501 | 170 | 70 x 60 x 40 | 2 | 1,22,800 | 1,63,100 | 1,70,300 | 5,000 |
| TI-502 | 200 | 60 x 60 x 60 | 2 | 1,27,600 | 1,70,300 | 1,78,200 | 5,500 |
| TI-503 | 285 | 80 x 60 x 60 | 3 | 1,44,200 | 1,89,000 | 1,98,200 | 5,500 |
| TI-504 | 325 | 90 x 60 x 60 | 4 | 1,53,600 | 2,06,700 | 2,16,100 | 6,000 |
| TI-505 | 400 | 110 x 60 x 60 | 4 | 1,79,400 | 2,37,900 | 2,46,100 | 8,000 |
| TI-506 | 600 | 125 x 70 x 70 | 4 | 2,15,400 | 2,77,800 | 2,91,100 | 10,000 |
| TI-507 | 800 | 125 x 80 x 80 | 5 | 2,46,200 | 3,09,600 | 3,28,500 | 10,800 |
| TI-508 | 1000 | 157 x 80 x 80 | 5 | 2,63,100 | 3,36,500 | 3,56,600 | 12,000 |
| TI-509 | 1500 Double Door | 146 x 150 x 69 | 10 | PRICE ON REQUEST |  |  | 14,000 |
| TI-510 | 2000 Double Door | 146 x 150 x 92 | 10 | PRICE ON REQUEST |  |  | 15,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

OUTER SIZE : TI – 509   225 H x 170 W x 130 D in cm

OUTER SIZE : TI – 510   225 H x 170 W x 150 D in cm

ERECTION CHARGES WILL BE EXTRA RS. 50,000/- FOR  (TI – 509 & 510 models are BIG & assembled type on SITE)

Note	               :  Two blowers for 400 ltrs to 2000 ltrs. Capacity chambers.

Validation & Documentation :  2 hrs. Loaded OR unloaded condition on one cycle. (4 or 8 temp. probes).

Along with IQ, OQ & PQ Documents.

Cost of Validation	               :  EXTRA Rs.  15,000/- + GST @ 18% Extra.

OPTIONALS :-

| A] | Complete PLC system (with touch screen HMI) :- |  |
| --- | --- | --- |
|  | PLC for auto changeover for standby refrigeration system with 7” Touch screen color HMI. | 45,000 |
|  | 21 CFR Part 11 compliance Ethernet based communication software for data management with multi dropping monitor 8 chambers to one software. | 30,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 100 ltr. to 325 ltr | 40,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 400 ltr. to 600 ltr | 50,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 800 ltr. to 2000 ltr | 62,000 |
|  | Scanner: Scanner 4 point temperature scanner complete with sensors. | 15,000 |
|  | Scanner: Scanner 8 point temperature scanner complete with sensors. | 38,000 |
|  | Mobile alert system. (Only for 5 nos. Mobile device) To be ordered with PLC System | 12,000 |
|  | Security System (Door Access) To be ordered with PLC System | 16,500 |
| B] | DATA LOGGER SYSTEM :- |  |
|  | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. | 22,000 |
|  | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. | 20,000 |
| C] | SCANNER :- |  |
| 1 | Scanner: Scanner 4 OR 8 point temperature scanner complete with sensors. | 33,000 |
| D] | OTHERS :- |  |
| 1 | Safety digital temperature controller with accessories | 6,000 |
| 2 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 3 | Stainless steel rod trays: 100 ltr. to 400 ltr. | 3,000 |
| 4 | Stainless steel rod trays: 600 ltr. to 2000 ltr. | 4,500 |

Requirements for Installation :-

- PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

- Place the chamber where ventilation is good

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 300 mm space from the sides and leave 300 mm space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V 5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

LABORATORY REFRIGERATOR (Cooling Incubator) :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8418 |
| Certification | CE Certification |
| Construction | Double walled construction with backside triple. |
| Insulation | 3” Thick PUF Insulation ensures stable temperature & reduced energy consumption. |
| Air Circulation | Motorized Blower at back side of the chamber develops unique air flow system which maximum uniform temperature condition inside the chamber. Unique air flow assures quick recovery after door openings. |
| Observation Door | Full view observation Acrylic door / Glass Door with gasket to observe sample inside the chamber. |
| Heating System | High quality S.S. Tubular Heaters are used for better heating. |
| Refrigeration system | CFC Free hermitically sealed Emerson Copeland make compressor for better cooling with R 134 A (Eco friendly) refrigerant with time delay to safe guard compressor system. |
| Validation Port | 50 mm validation port hole with Silicon rubber seal to insert sensors for validation. |
| Lighting | Interior illumination for Working area. |
| Safety features | - In case of Temperature variation, Temperature deviation alarm will active & give audio visual alarm. - Safety Thermostat Cut off the supply in case temperature goes above 350 C & activate safety alarm. - Compressor ON delay timer (2 min.) to safe guard the compressor. - In built over load protector provided for hermetically sealed compressor. |
| Features | - Fiber/PU Wheel with front lockable for easy movement. - MCB for electrical safety - Adjustable tray height arrangements. |
| Trays | Specially designed stainless steel rod trays ensure uniform temperature distribution. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 & Outer Galvanized Iron (GI) sheet powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304 |
| G.M.P. Model | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range | 20C to 80C |
| Temperature Accuracy | ± 0.50 C |
| Temperature Uniformity | ± 20 C |
| Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value. |

| Models | Capacity Liters | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | G.M.P. Model (Amt. in Rs.) | Packing Charge |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TI-810 | 100 | 50 x 50 x 40 | 2 | 1,07,800 | 1,36,200 | 1,42,100 | 5,000 |
| TI-811 | 170 | 70 x 60 x 40 | 2 | 1,22,900 | 1,63,300 | 1,70,300 | 5,000 |
| TI-812 | 200 | 60 x 60 x 60 | 2 | 1,27,600 | 1,70,300 | 1,78,200 | 5,500 |
| TI-813 | 285 | 80 x 60 x 60 | 3 | 1,44,200 | 1,86,900 | 1,98,100 | 5,500 |
| TI-814 | 325 | 90 x 60 x 60 | 4 | 1,53,600 | 2,06,800 | 2,16,100 | 6,000 |
| TI-815 | 400 | 110 x 60 x 60 | 4 | 1,79,400 | 2,37,900 | 2,46,100 | 8,000 |
| TI-816 | 600 | 125 x 70 x 70 | 4 | 2,15,400 | 2,77,800 | 2,91,100 | 10,000 |
| TI-817 | 800 | 125 x 80 x 80 | 5 | 2,46,200 | 3,09,600 | 3,28,500 | 10,800 |
| TI-818 | 1000 | 157 x 80 x 80 | 5 | 2,63,100 | 3,36,400 | 3,56,500 | 12,000 |
| TI-819 | 1500 Double Door | 146 x 150 x 69 | 10 | PRICE ON REQUEST |  |  | 14,000 |
| TI-820 | 2000 Double Door | 146 x 150 x 92 | 10 | PRICE ON REQUEST |  |  | 15,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

OUTER SIZE : TI – 819   225 H x 170 W x 130 D in cm

OUTER SIZE : TI – 820   225 H x 170 W x 150 D in cm

ERECTION CHARGES WILL BE EXTRA RS. 50,000/- FOR (TI – 819 & 820 models are BIG & assembled type on SITE)

Note	               :  Two blowers for 400 ltrs to 2000 ltrs. Capacity chambers.

Validation & Documentation :  2 hrs. Loaded OR unloaded condition on one cycle. (4 or 8 temp. probes).

Along with IQ, OQ & PQ Documents.

Cost of Validation	               :  EXTRA Rs.  15,000/- + GST @ 18% Extra.

OPTIONALS :-

| A] | Complete PLC system (with touch screen HMI) :- |  |
| --- | --- | --- |
|  | PLC for auto changeover for standby refrigeration system with 7” Touch screen color HMI. | 45,000 |
|  | 21 CFR Part 11 compliance Ethernet based communication software for data management with multi dropping monitor 8 chambers to one software. | 30,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 100 ltr. to 325 ltr | 40,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 400 ltr. to 600 ltr | 50,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 800 ltr. to 2000 ltr | 62,000 |
|  | Scanner: Scanner 4 point temperature scanner complete with sensors. | 15,000 |
|  | Scanner: Scanner 8 point temperature scanner complete with sensors. | 38,000 |
|  | Mobile alert system. (Only for 5 nos. Mobile device) To be ordered with PLC System | 12,000 |
|  | Security System (Door Access) To be ordered with PLC System | 16,500 |
| B] | DATA LOGGER SYSTEM :- |  |
|  | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. | 22,000 |
|  | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. | 20,000 |
| C] | SCANNER :- |  |
| 1 | Scanner: Scanner 4 OR 8 point temperature scanner complete with sensors. | 33,000 |
| D] | OTHERS :- |  |
| 1 | Safety digital temperature controller with accessories | 6,000 |
| 2 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 3 | Stainless steel rod trays: 100 ltr. to 400 ltr. | 3,000 |
| 4 | Stainless steel rod trays: 600 ltr. to 2000 ltr. | 4,500 |

Requirements for Installation :-

- PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

- Place the chamber where ventilation is good

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 300 mm space from the sides and leave 300 mm space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V 5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

Glass Door Laboratory Refrigerator  :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8418 |
| Certification | CE Certification |
| Construction | Double walled construction with backside triple. |
| Insulation | 3” Thick PUF Insulation ensures stable temperature & reduced energy consumption. |
| Air Circulation | Motorized Blower at back side of the chamber develops unique air flow system which maximum uniform temperature condition inside the chamber. Unique air flow assures quick recovery after door openings. |
| Doors | Front door Multi Glass fitted in Mild steel Powder coated Or S. S. 304 with full length magnetic gasket for proper sealing with lock and key. Full view observation Glass door with gasket to observe sample inside the chamber. |
| Heating System | High quality S.S. Tubular Heaters are used for better heating. |
| Refrigeration system | CFC Free hermitically sealed Emerson Copeland make compressor for better cooling with R134A (Eco friendly) refrigerant with time delay to safe guard compressor system. |
| Validation Port | 50 mm validation port hole with Silicon rubber seal to insert sensors for validation. |
| Lighting | Interior illumination for Working area. |
| Safety features | - In case of Temperature variation, Temperature deviation alarm will active & give audio visual alarm. - Safety Thermostat Cutoff the supply in case temperature goes above 350C & activate safety alarm. - Compressor ON delay timer (2 min.) to safe guard the compressor. - In built over load protector provided for hermetically sealed compressor. |
| Features | - Fiber/PU Wheel with front lockable for easy movement. - MCB for electrical safety. - Adjustable tray height arrangements. |
| Trays | Specially designed stainless steel rod trays ensure uniform temperature distribution. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 & Outer Galvanized Iron (GI) sheet powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S. S. 304 |
| GMP Model | Inside S.S. 316 & Outer S. S. 304 |
| Temperature Range | 20C to 80C |
| Temperature Accuracy | ± 0.50 C |
| Temperature Uniformity | ± 20 C |
| Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value. |

| Models | Capacity Liters | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | G.M.P. Model (Amt. in Rs.) | Packing Charge |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TI-840 | 100 | 50 x 50 x 40 | 2 | 1,10,600 | 1,44,000 | 1,47,400 | 5,000 |
| TI-841 | 170 | 70 x 60 x 40 | 2 | 1,27,300 | 1,75,400 | 1,80,000 | 5,000 |
| TI-842 | 200 | 60 x 60 x 60 | 2 | 1,32,000 | 1,83,100 | 1,88,400 | 5,500 |
| TI-843 | 285 | 80 x 60 x 60 | 3 | 1,50,300 | 2,08,900 | 2,14,400 | 5,500 |
| TI-844 | 325 | 90 x 60 x 60 | 4 | 1,58,100 | 2,19,900 | 2,26,100 | 6,000 |
| TI-845 | 400 | 110 x 60 x 60 | 4 | 1,83,700 | 2,48,300 | 2,55,500 | 8,000 |
| TI-846 | 600 | 125 x 70 x 70 | 4 | 2,18,100 | 2,83,700 | 2,94,500 | 10,000 |
| TI-847 | 800 | 125 x 80 x 80 | 5 | 2,49,300 | 3,21,300 | 3,32,500 | 10,800 |
| TI-848 | 1000 | 157 x 80 x 80 | 5 | 2,67,100 | 3,50,700 | 3,63,000 | 12,000 |
| TI-849 | 1500 Double Door | 146 x 150 x 69 | 10 | PRICE ON REQUEST |  |  | 14,000 |
| TI-850 | 2000 Double Door | 146 x 150 x 92 | 10 | PRICE ON REQUEST |  |  | 15,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

OUTER SIZE : TI – 849   225 H x 170 W x 130 D in cm

OUTER SIZE : TI – 850   225 H x 170 W x 150 D in cm

ERECTION CHARGES WILL BE EXTRA RS.50,000/- (TI – 849 & 850 models are BIG & assembled type on SITE)

Note	               :  Two blowers for 400 ltrs to 2000 ltrs. Capacity chambers.

Validation & Documentation :  2 hrs. Loaded OR unloaded condition on one cycle. (4 or 8 temp. probes).

Along with IQ, OQ & PQ Documents.

Cost of Validation	               :  EXTRA Rs.  15,000/- + GST @ 18% Extra.

OPTIONALS :-

| A] | Complete PLC system (with touch screen HMI) :- |  |
| --- | --- | --- |
|  | PLC for auto changeover for standby refrigeration system with 7” Touch screen color HMI. | 45,000 |
|  | 21 CFR Part 11 compliance Ethernet based communication software for data management with multi dropping monitor 8 chambers to one software. | 30,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 100 ltr. to 325 ltr | 40,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 400 ltr. to 600 ltr | 50,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 800 ltr. to 2000 ltr | 62,000 |
|  | Scanner: Scanner 4 point temperature scanner complete with sensors. | 15,000 |
|  | Scanner: Scanner 8 point temperature scanner complete with sensors. | 38,000 |
|  | Mobile alert system. (Only for 5 nos. Mobile device) To be ordered with PLC System | 12,000 |
|  | Security System (Door Access) To be ordered with PLC System | 16,500 |
| B] | DATA LOGGER SYSTEM :- |  |
|  | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. | 22,000 |
|  | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. | 20,000 |
| C] | SCANNER :- |  |
| 1 | Scanner: Scanner 4 OR 8 point temperature scanner complete with sensors. | 33,000 |
| D] | OTHERS :- |  |
| 1 | Safety digital temperature controller with accessories | 6,000 |
| 2 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 3 | Stainless steel rod trays: 100 ltr. to 400 ltr. | 3,000 |
| 4 | Stainless steel rod trays: 600 ltr. to 2000 ltr. | 4,500 |

Requirements for Installation :-

- PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

- Place the chamber where ventilation is good

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 300 mm space from the sides and leave 300 mm space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V 5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

DEEP FREEZER (-200C) :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8418 |
| Certification | CE Certification |
| Construction | Double walled construction. |
| Insulation | 3” Thick PUF Insulation ensures temperature stability with reduced energy consumption. |
| Inner Door | PUF Insulated Stainless Steel inner door with gasket. |
| Refrigeration system | By means of CFC Free hermitically sealed Emerson Copland make compressor with R 404 A for better cooling with ECO friendly refrigerant with time delay to safe guard compressor system. |
| Safety | - Compressor ON delay timer (2 min.) to safe guard the compressor. - In built over load protector provided for hermetically sealed compressor. |
| Features | - Fiber/PU Wheel with front lockable for easy movement. - MCB for electrical safety. - Adjustable tray height arrangements. |
| Trays | Specially designed stainless steel rod trays ensure uniform temperature distribution. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 & Outer Galvanized Iron (GI) sheet powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304 |
| G.M.P Model | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range | - 100 C to - 200 C |
| Temperature Accuracy | ± 30 C |
| Temperature Uniformity | + 50 C |
| Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value. |

| Models | Cap. Liters | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | G.M.P. Model (Amt. in Rs.) | Packing Charge |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TI-821 | 100 | 50 x 50 x 40 | 2 | 1,09,700 | 1,46,500 | 1,47,200 | 5,000 |
| TI-822 | 170 | 70 x 60 x 40 | 2 | 1,23,400 | 1,69,300 | 1,70,200 | 5,000 |
| TI-823 | 200 | 60 x 60 x 60 | 2 | 1,31,000 | 1,73,000 | 1,80,300 | 5,500 |
| TI-824 | 285 | 80 x 60 x 60 | 3 | 1,54,600 | 2,06,800 | 2,08,700 | 5,500 |
| TI-825 | 325 | 90 x 60 x 60 | 4 | 1,64,900 | 2,23,000 | 2,25,400 | 6,000 |
| TI-826 | 400 | 110 x 60 x 60 | 4 | 1,94,100 | 2,61,800 | 2,64,200 | 8,000 |
| TI-827 | 500 | 105 x 70 x 70 | 4 | 2,04,600 | 2,76,300 | 2,80,200 | 10,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

Validation & Documentation :  2 hrs. Loaded OR unloaded condition on one cycle. (4 or 8 temp. probes).

Along with IQ, OQ & PQ Documents.

Cost of Validation	               :  EXTRA Rs.  15,000/- + GST @ 18% Extra.

OPTIONALS :-

| A] | Complete PLC system (with touch screen HMI) :- |  |
| --- | --- | --- |
|  | PLC for auto changeover for standby refrigeration system with 7” Touch screen color HMI. | 55,000 |
|  | 21 CFR Part 11 compliance Ethernet based communication software for data management with multi dropping monitor 8 chambers to one software. | 30,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 100 ltr. to 200 ltr | 50,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 285 ltr. to 500 ltr | 70,000 |
|  | Scanner: Scanner 4 point temperature scanner complete with sensors. | 15,000 |
|  | Scanner: Scanner 8 point temperature scanner complete with sensors. | 38,000 |
|  | Mobile alert system. (Only for 5 nos. Mobile device) To be ordered with PLC System | 12,000 |
|  | Security System (Door Access) To be ordered with PLC System | 16,500 |
| B] | DATA LOGGER SYSTEM :- |  |
|  | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. | 22,000 |
|  | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. | 20,000 |
| C] | SCANNER :- |  |
| 1 | Scanner: Scanner 4 OR 8 point temperature scanner complete with sensors. | 33,000 |
| D] | OTHERS :- |  |
| 1 | Safety digital temperature controller with accessories | 6,000 |
| 2 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 3 | Stainless steel rod trays: 100 ltr. to 500 ltr. | 3,000 |

Requirements for Installation :-

- PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

- Place the chamber where ventilation is good

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 300 mm space from the sides and leave 300 mm space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V 5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

DEEP FREEZER (- 400C)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8418 |
| Certification | CE Certification |
| Construction | Double walled construction |
| Insulation | 4” Thick PUF Insulation ensures temperature stability & reduces energy consumption. |
| Inner Door | PUF Insulated Stainless Steel inner door with gasket. |
| Refrigeration system | By means of CFC Free hermitically sealed Emerson Copland make compressor with R 404 A for better cooling with ECO friendly refrigerant with time delay to safe guard compressor system. |
| Safety | - Compressor ON delay timer (2 min.) to safe guard the compressor. - In built over load protector provided for hermetically sealed compressor. |
| Features | - Fiber/PU Wheel with front lockable for easy movement. - MCB for electrical safety. - Adjustable tray height arrangements. |
| Trays | Specially designed stainless steel Rod trays ensure uniform temperature distribution. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 & Outer Galvanized Iron (GI) sheet powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304 |
| GMP Model | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range | -300C to -400C |
| Temperature Accuracy | ± 30C |
| Temperature Uniformity | + 50 C |
| Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value. |

| Models | Cap. Liters | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | G.M.P. Model (Amt. in Rs.) | Packing Charge |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TI-831 | 100 | 50 x 50 x 40 | 2 | 1,16,500 | 1,54,600 | 1,58,800 | 5,000 |
| TI-832 | 170 | 70 x 60 x 40 | 2 | 1,36,700 | 1,83,600 | 1,89,300 | 5,000 |
| TI-833 | 200 | 60 x 60 x 60 | 2 | 1,43,500 | 1,93,200 | 2,00,200 | 5,500 |
| TI-834 | 285 | 80 x 60 x 60 | 3 | 1,78,300 | 2,31,900 | 2,40,900 | 5,500 |
| TI-835 | 325 | 90 x 60 x 60 | 4 | 1,86,600 | 2,45,700 | 2,56,000 | 6,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

Validation & Documentation :  2 hrs. Loaded OR unloaded condition on one cycle. (4 or 8 temp. probes).

Along with IQ, OQ & PQ Documents.

Cost of Validation	               :  EXTRA Rs.15,000/- + GST @ 18% Extra.

OPTIONALS :-

| A] | Complete PLC system (with touch screen HMI) :- |  |
| --- | --- | --- |
|  | PLC for auto changeover for standby refrigeration system with 7” Touch screen color HMI. | 55,000 |
|  | 21 CFR Part 11 compliance Ethernet based communication software for data management with multi dropping monitor 8 chambers to one software. | 30,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 100 ltr. to 285 ltr | 63,000 |
|  | Ready to use stand by refrigeration system in case of regular system fails. Complete with Compressor, Condenser, Cooling coil etc. : 325 ltr | 85,000 |
|  | Scanner: Scanner 4 point temperature scanner complete with sensors. | 15,000 |
|  | Scanner: Scanner 8 point temperature scanner complete with sensors. | 38,000 |
|  | Mobile alert system. (Only for 5 nos. Mobile device) To be ordered with PLC System | 12,000 |
|  | Security System (Door Access) To be ordered with PLC System | 16,500 |
| B] | DATA LOGGER SYSTEM :- |  |
|  | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. | 22,000 |
|  | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. | 20,000 |
| C] | SCANNER :- |  |
| 1 | Scanner: Scanner 4 OR 8 point temperature scanner complete with sensors. | 33,000 |
| D] | OTHERS :- |  |
| 1 | Safety digital temperature controller with accessories | 6,000 |
| 2 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 3 | Stainless steel rod trays: 100 ltr. to 500 ltr. | 3,000 |

Requirements for Installation :-

- PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 50mm space from the sides and leave 300mm space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V + 5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

- Stabilized: 230V AC single-phase supply. MCB with proper Earthing.

PRODUCT RANGE (HEATING)

LABORATORY OVEN   (Natural Convection): 2500 C

(NOTE : DO NOT USE FOR CHEMICAL DRYING)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Construction | Double walled construction. |
| Insulation | 3” High grade glass wool insulation. |
| Heating System | Heating elements made from the best quality Kanthal DSD wire are used on refractory supports. |
| Safety | Safety cum stand by thermostat to cut off the Heater supply in case of overshoot of temperature. |
| Door | Single door with asbestos gasket, fitted on heavy cast and chromium plated brass hinges with spring and roller type latch. |
| Trays | Adjustable Stainless steel Rod trays. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 & Outer Mild steel (CRCA) sheet with powder coated. |
| Temperature Range | 50C above ambient to 2500C |
| Temperature Accuracy | ± 0.50C |
| Control System | Imported microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |

| Models | Capacity Ltrs. | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-125 A | 43 | 35 x 35 x 35 | 2 | 40,800 | 2,000 |
| TI-125 B | 91 | 45 x 45 x 45 | 2 | 48,900 | 3,000 |
| TI-125 C | 122 | 60 x 45 x 45 | 3 | 56,500 | 3,500 |
| TI-125 D | 216 | 60 x 60 x 60 | 3 | 69,400 | 4,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.`

Requirement for Installation :

- Place the chamber where ventilation is good.  The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 300 mm space from the sides and leave 300 mm space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

LABORATORY DRYING OVEN 250 0 C  :-

(NOTE : DO NOT USE FOR CHEMICAL DRYING)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Construction | Double walled construction. |
| Insulation | 3” High grade glass wool insulation. |
| Air Draft | Forced air draft is provided by means of Blower fitted at the bottom of the Oven (There is no re-circulation of air). |
| Blower Motor | 1/12 HP, TEFC , F-class insulation, Single Phase, 1440 RPM, 230 Volts |
| Heating System | Heating elements made from the best quality Kanthal DSD wire are used on refractory supports. |
| Safety | Safety cum stand by thermostat to cut off the Heater supply in case of overshoot of temperature. |
| Door | Single door fitted on heavy cast and chromium plated brass hinges with spring and roller type latch. |
| Trays | Adjustable Stainless steel Rod trays. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 & Outer Mild steel (CRCA) sheet with powder coated. |
| Temperature Range | 500C to 2500 C |
| Temperature Accuracy | ± 0.50 C |
| Control System | Imported microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |

| Models | Capacity Liters | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-126 A | 43 | 35 x 35 x 35 | 2 | 53,800 | 2,500 |
| TI-126 B | 91 | 45 x 45 x 45 | 2 | 63,300 | 3,000 |
| TI-126 C | 122 | 60 x 45 x 45 | 3 | 71,800 | 3,600 |
| TI-126 D | 216 | 60 x 60 x 60 | 3 | 86,100 | 4,100 |
| TI-126 E | 324 | 90 x 60 x 60 | 4 | 1,02,300 | 5,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

Requirement for Installation :

- Place the chamber where ventilation is good.  The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 300 mm space from the sides and leave 300 mm space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

LABORATORY PRECISION OVEN :    Uniform Temperature  2500 C

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Construction | Tripled walled construction. |
| Insulation | 3” High grade heavy glass wool insulation ensures stable temperature & reduced energy consumption. |
| Air Recirculation | Motorized Blower on top side of the chamber develops unique air flow system which ensures maximum uniform temperature distribution inside the chamber. Unique air flow assures quick recovery after door openings. |
| Blower Motor | 1/12 HP, TEFC , F-class insulation, Single Phase, 1440 RPM, 230 Volts |
| Heating System | S. S. tubular heating elements used on side walls to maintain uniform condition inside the chamber. |
| Safety | Safety cum stand by thermostat to cut off the heater supply in case of overshoot of temperature. |
| Door | Standard Model : Single door fitted on heavy cast and chromium plated hinges with spring and roller type latch with Asbestos gasket. (NOTE : Roller type latch for TI-128 A to TI-128 D Toggel Clamp Lock for TI-128 E to TI-128 I) Inner & Outer SS 304 / GMP Model : Single door fitted on heavy cast and chromium plated hinges with Toggel Clamp Lock with Silicon gasket. |
| Trays | Specially designed Stainless Steel Rod Trays which ensure uniform temp. distribution. |
| Supply | 230V AC, Single Phase, 50Hz (Model TI-128 A to 128 E) 415V AC, Three Phase, 50 Hz (Model TI-128 F to TI-128 I) |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 & Outer Mild steel (CRCA) sheet with powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304 |
| G.M.P. Model | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range | 50 C above ambient to 2500 C : For Standard Model 50 C above ambient to 2500 C : For Inner & Outer SS 304 / G.M.P. Model NOTE : Especially in a Inner & Outer SS 304 / G.M.P. Model Outer is made of Stainless Steel, it is not possible to achieve low skin temperature as compare to STD Model which have heat resistant thick paint applied on it. FRONT DOOR : 95 to 1000 C (Approx.) on 2250C Set Point. OUTER BOTH SIDE : 70 to 800 C (Approx.) |
| Temperature Accuracy | ± 0.20 C |
| Temperature Uniformity | + 50 C |
| Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |

| Models | Capacity Liters | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | G.M.P. Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Following Models are Single Phase 230V AC |  |  |  |  |  |  |  |
| TI-128 A | 43 | 35 x 35 x 35 | 2 | 53,800 | 77,100 | 83,000 | 2,500 |
| TI-128 B | 91 | 45 x 45 x 45 | 2 | 63,900 | 92,700 | 99,900 | 3,300 |
| TI-128 C | 122 | 60 x 45 x 45 | 3 | 72,100 | 1,06,600 | 1,16,900 | 3,700 |
| TI-128 D | 216 | 60 x 60 x 60 | 3 | 82,700 | 1,22,800 | 1,35,800 | 3,700 |
| TI-128 E | 324 | 90 x 60 x 60 | 4 | 1,07,700 | 1,54,500 | 1,73,700 | 5,000 |
| Following Models are Three Phase 415V AC |  |  |  |  |  |  |  |
| TI-128 F | 400 | 110 x 60 x 60 | 4 | 1,43,300 | 1,95,700 | 2,23,000 | 7,500 |
| TI-128 G | 600 | 125 x 70 x 70 | 4 | 1,81,500 | 2,39,700 | 2,77,400 | 9,000 |
| TI-128 H | 800 | 125 x 80 x 80 | 5 | 1,98,500 | 2,69,200 | 3,11,200 | 11,000 |
| TI-128 I | 1000 | 157 x 80 x 80 | 5 | 2,17,300 | 2,98,400 | 3,44,500 | 15,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

STANDARD MODEL WITH SILICON GASKET WILL BE CHARGE EXTRA.

Validation & Documentation :  2 hrs. Loaded OR unloaded condition on one cycle. (4 or 8 temp. probes).

Along with IQ, OQ & PQ Documents.

Cost of Validation	               :  EXTRA Rs.  15,000/- + GST @ 18% Extra.

NOTE :- Packing Charges – In case of unpacked condition (without wooden box packing) Pallet charges will be extra Rs. 3,000/- for TI-128 G to TI-128 I

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Scanner : Scanner 4 or 8 point temperature scanner complete with sensors. | 33,000 |
| 4 | Safety digital temperature controller with accessories | 6,000 |
| 5 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 6 | Stainless steel rod trays : 100 ltr. to 324 ltr. | 3,000 |
| 7 | Stainless steel rod trays : 400 ltr. to 1000 ltr. | 4,500 |
| 8 | Profile controller / Programmable controller - 8 ramp & 8 sack period | 8,000 |
| 9 | Motor speed controller for Air flow control application. (Between 1000 to 1400 RPM) | 4,500 |

Requirements for Installation :-

- PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port

- Place the chamber where ventilation is good.  The chamber must be operated in an ambient of 220C ~ 300C

- Leave minimum 50mm space from the sides and leave 300mm space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V + 5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures

LABORATORY PRECISION OVEN    Uniform Temperature 300 0 C  :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Construction | Tripled walled construction. |
| Insulation | 4” High grade heavy Glass wool insulation ensures stable temperature & reduced energy consumption. |
| Air Recirculation | Motorized Blower with insulated shaft on back side of the chamber develops unique air flow system which ensures maximum uniform temperature distribution inside the chamber. Unique air flow assures quick recovery after door openings. |
| Blower Motor | 1/12 HP, TEFC , F-class insulation, Single Phase, 1440 RPM, 230 Volts |
| Heating System | S. S. tubular heating elements used on back walls to maintain uniform condition inside the chamber. |
| Safety PID Controller | Safety PID controller to cut-off the Heater supply in case of temperature overshoot with Alarm System. |
| Door | Standard Model : Single door fitted on heavy cast and chromium plated hinges with spring and roller type latch with ceramic gasket. (NOTE : Roller type latch for TI-130 A to TI-130 D Toggel Clamp Lock for TI-130 E Inner & Outer SS 304 Model : Single door fitted on heavy cast and chromium plated hinges with Toggel Clamp locking with Viton gasket GMP Model : Single door fitted on heavy cast and chromium plated hinges with Toggel Clamp locking with Viton gasket |
| Trays | Specially designed stainless steel rod trays ensure uniform temperature distribution. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 & Outer Mild steel (CRCA) sheet with powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304 |
| G.M.P. Model | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range | 50 C above ambient to 3000 C NOTE : Especially in a Inner & Outer SS 304 / G.M.P. Model Outer is made of Stainless Steel, it is not possible to achieve low skin temperature as compare to STD Model which have heat resistant thick paint applied on it. FRONT DOOR : 95 to 1000 C (Approx.) on 3000C Set Point. BOTH SIDES : 70 to 800 C (Approx.) |
| Temperature Accuracy | ± 0.20 C |
| Temperature Uniformity | + 50 C |
| Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |

| Models | Cap. Liters | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | G.M.P. Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TI-130 A | 43 | 35 x 35 x 35 | 2 | 62,100 | 91,900 | 98,200 | 2,500 |
| TI-130 B | 91 | 45 x 45 x 45 | 2 | 70,900 | 1,09,100 | 1,18,900 | 3,300 |
| TI-130 C | 122 | 60 x 45 x 45 | 3 | 80,000 | 1,23,300 | 1,35,800 | 3,700 |
| TI-130 D | 216 | 60 x 60 x 60 | 3 | 94,400 | 1,43,900 | 1,61,400 | 3,700 |
| TI-130 E | 324 | 90 x 60 x 60 | 4 | 1,17,000 | 1,77,600 | 2,01,700 | 5,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

STANDARD MODEL WITH VITON GASKET; WILL BE CHARGE EXTRA.

Validation & Documentation :  2 hrs. Loaded OR unloaded condition on one cycle. (4 or 8 temp. probes).

Along with IQ, OQ & PQ Documents.

Cost of Validation	               :  EXTRA Rs.  15,000/- + GST @ 18% Extra.

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Scanner : Scanner 4 or 8 point temperature scanner complete with sensors. | 33,000 |
| 4 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 5 | Stainless steel rod trays : 100 ltr. To 324 ltr. | 3,000 |
| 6 | Profile controller / Programmable controller – 8 ramp & 8 sack period | 8,000 |

Requirements for Installation :-

            - PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

            - Place the chamber where ventilation is good.

            - The chamber must be operated in an ambient of 220C ~ 300C.

            - Leave minimum 50mm space from the sides and leave 300mm space from the backside.

            - Place the chamber horizontally evelled.

            - Give stabilized electric supply (230V + 5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use  voltage stabilizer (KVA).

            - Proper grounding of the chamber to get satisfactory operation of chamber.

            - Do not test any product in the chamber, which is capable of generating combustible mixtures.

LABORATORY PRECISION OVEN With Glass Window On Door : Uniform Temperature  2500 C

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Construction | Tripled walled construction. |
| Insulation | 3” High grade heavy glass wool insulation ensures stable temperature & reduced energy consumption. |
| Air Recirculation | Motorized Blower on top side of the chamber develops unique air flow system which ensures maximum uniform temperature distribution inside the chamber. Unique air flow assures quick recovery after door openings. |
| Blower Motor | 1/12 HP, TEFC , F-class insulation, Single Phase, 1440 RPM, 230 Volts |
| Heating System | S. S. tubular heating elements used on side walls to maintain uniform condition inside the chamber. |
| Safety | Safety cum stand by thermostat to cut off the heater supply in case of overshoot of temperature. |
| Door With Glass Window | Standard Model : Single door fitted on heavy cast and chromium plated hinges with spring and roller type latch with Asbestos gasket. (NOTE : Roller type latch for TI-131 A to TI-131 D Toggel Clamp Lock for TI-131 E Inner & Outer SS 304 Model : Single door fitted on heavy cast and chromium plated hinges with Toggel Clamp Lock with Silicon gasket G.M.P. Model : Single door fitted on heavy cast and chromium plated hinges with Toggel Clamp Lock with Silicon gasket |
| Glass Window on Door | Door construction having double tempered glass window within the main door. |
| Trays | Specially designed Stainless Steel Rod Trays which ensure uniform temp.distribution. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 Outer Mild steel (CRCA) sheet with powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 Outer S.S. 304 |
| G.M.P. Model | Inside S.S. 316 Outer S.S. 304 |
| Temperature Range | 50 C above ambient to 2500 C : For Standard Model 50 C above ambient to 2250 C : For Inner & Outer SS 304 / G.M.P. Model NOTE : Especially in a Inner & Outer SS 304 / G.M.P. Model, Outer is made of Stainless Steel, it is not possible to achieve low skin temperature as compare to STD Model which have heat resistant thick paint applied on it. FRONT DOOR : 95 to 1000 C (Approx.) on 2250C Set Point. BOTH SIDE : 70 to 800 C (Approx.) |
| Temperature Accuracy | ± 0.20 C |
| Temperature Uniformity | + 50 C |
| Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |

| Models | Cap. Ltr. | Internal Size H x W x D in cm | Trays | Glass Window (Size in cm) | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | G.M.P. Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| TI-131 A | 43 | 35 x 35 x 35 | 2 | 20 x 20 | 57,500 | 83,100 | 87,600 | 2,500 |
| TI-131 B | 91 | 45 x 45 x 45 | 2 | 20 x 20 | 69,400 | 1,00,800 | 1,06,300 | 3,300 |
| TI-131 C | 122 | 60 x 45 x 45 | 3 | 30 x 30 | 78,700 | 1,16,300 | 1,23,700 | 3,700 |
| TI-131 D | 216 | 60 x 60 x 60 | 3 | 30 x 30 | 91,600 | 1,33,900 | 1,56,400 | 3,700 |
| TI-131 E | 324 | 90 x 60 x 60 | 4 | 45 x 30 | 1,18,200 | 1,69,900 | 2,00,000 | 5,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

STANDARD MODEL WITH SILICON GASKET WILL BE CHARGE EXTRA.

Validation & Documentation :  2 hrs. Loaded OR unloaded condition on one cycle. (4 or 8 temp. probes).

Along with IQ, OQ & PQ Documents.

Cost of Validation	               :  EXTRA Rs.  15,000/- + GST @ 18% Extra.

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Scanner : Scanner 4 or 8 point temperature scanner complete with sensors. | 33,000 |
| 4 | Safety digital temperature controller with accessories | 6,000 |
| 5 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 6 | Stainless steel rod trays : 100 ltr. To 324 ltr. | 4,000 |
| 7 | Profile controller / Programmable controller – 8 ramp & 8 sack period | 8,000 |
| 8 | Motor speed controller for Air flow control application. | 4,500 |

Requirements for Installation :-

      - PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

      - Place the chamber where ventilation is good.The chamber must be operated in an ambient of 220C~300C.

      - Leave minimum 50mm space from the sides and leave 300mm space from the backside.

      - Place the chamber horizontally evelled.

      - Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a voltage stabilizer (KVA).

      - Proper grounding of the chamber to get satisfactory operation of chamber.

      - Do not test any product in the chamber, which is capable of generating combustible mixtures

LABORATORY PRECISION OVEN  High Temperature  400 0C  :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Construction | Tripled walled construction with Outer Hammer Tone Paint. |
| Insulation | 5” High grade heavy glass wool insulation ensures stable temperature & reduced energy consumption. |
| Air Recirculation ( V-Belt Drive ) | Motorized Blower with V-BELT drive on top side of the chamber develops unique air flow system which ensures maximum uniform temp. distribution inside the chamber. Unique air flow assures quick recovery after door openings. V-Belt drive ensures continuous run of motor on high temperature to avoid damage of motor. |
| Blower Motor | 1/4 HP, TEFC , F-class insulation, Single Phase, 1440 RPM, 230 Volts |
| Heating System | Specially designed U shape S. S. tubular heating elements used on side walls to maintain uniform condition inside the chamber. Heaters terminals taken out for wiring to avoid burning of heater terminals on high temperature. |
| Safety PID Controller | - Safety PID controller to cut-off the Heater supply in case of temperature overshoot with Alarm System. - Limit switch for interlocking connection between Motor & Door. i.e. when door is open; motor will be switch off. |
| Door | Single door fitted on heavy cast and chromium plated hinges with Toggle Clamp lock with Ceramic gasket |
| Trays | Specially designed stainless steel wire mesh trays ensure uniform temperature distribution. |
| Supply | 230V AC, Single Phase, 50Hz (Model TI-128 A HT to 128 B HT) 415V AC, Three Phase, 50 Hz (Model TI-128 C HT to TI-128 E HT) |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 & Outer Mild steel (CRCA) sheet with hammer tone paint. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304 |
| G.M.P. Model | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range | 500C to 4000C FRONT DOOR : 95 to 1000 C (Approx.) on 4000C Set Point. BOTH SIDE : 70 to 800 C (Approx.) NOTE : Especially in a Inner & Outer SS 304 / G.M.P. Model, Outer is made of Stainless Steel, it is not possible to achieve low skin temperature as compare to STD Model which have heat resistant thick paint applied on it. |
| Temperature Accuracy | ± 0.50 C |
| Temperature Uniformity | + 100 C |
| Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |

| Models | Cap. Liters | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | G.M.P. Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Following Models are Single Phase 230V AC |  |  |  |  |  |  |  |
| TI-128 A HT | 43 | 35 x 35 x 35 | 2 | 98,500 | 1,23,000 | 1,35,600 | 4,500 |
| TI-128 B HT | 91 | 45 x 45 x 45 | 2 | 1,10,600 | 1,42,000 | 1,57,800 | 4,500 |
| Following Models are Three Phase 415V AC |  |  |  |  |  |  |  |
| TI-128 C HT | 122 | 60 x 45 x 45 | 3 | 1,24,400 | 1,57,900 | 1,75,000 | 5,500 |
| TI-128 D HT | 216 | 60 x 60 x 60 | 3 | 1,44,400 | 1,82,300 | 2,04,500 | 6,500 |
| TI-128 E HT | 324 | 90 x 60 x 60 | 4 | 1,75,800 | 2,22,600 | 2,53,800 | 10,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

NOTE: WHEN HIGH TEMPERATURE OVEN IS USE ABOVE 3000C THEN BLACK SPOTS OR STAINS GENERATED ON INNER CHAMBER, TO AVOID THE SAME WE APPLY HEAT RESISTANCE SILVER PAINT ON INNER CHAMBER.

Validation & Documentation :  2 hrs. Loaded OR unloaded condition on one cycle. (4 or 8 temp. probes).

Along with IQ, OQ & PQ Documents.

Cost of Validation	               :  EXTRA Rs.  15,000/- + GST @ 18% Extra.

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 4 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm. | 4,000 |
| 5 | Stainless steel rod trays : (43 ltr. To 122 ltr.) | 4,000 |
| 6 | Stainless steel rod trays : (216 ltr. To 324 ltr.) | 5,500 |
| 7 | Profile controller / Programmable controller – 8 ramp & 8 sack period | 8,000 |

Requirements for Installation :-

            - PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

            - Place the chamber where ventilation is good.

            - The chamber must be operated in an ambient of 220 C ~ 300C.

            - Leave minimum 50 mm space from the sides and leave 300mm space from the backside.

            - Place the chamber horizontally evelled.

            - Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 32 Amp).It is recommended to use a voltage stabilizer (KVA).

            - Proper grounding of the chamber to get satisfactory operation of chamber.

            - Do not test any product in the chamber, which is capable of generating combustible mixtures.

HIGH TEMPERATURE LABORATORY PRECISION OVEN  450 0C :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Construction | Tripled walled construction with Outer Hammer Tone Paint. |
| Insulation | 6” High grade heavy Cerewool insulation ensures stable temperature & reduced energy consumption. |
| Air Recirculation ( V-Belt Drive ) | - Motorized Blower with V-BELT drive on top side of the chamber develops unique air flow system which ensures maximum uniform temp. distribution inside the chamber. - Unique air flow assures quick recovery after door openings. - V-Belt drive ensures continuous run of motor on high temperature to avoid damage of motor. |
| Blower Motor | 1/4 HP, TEFC , F-class insulation, Single Phase, 1440 RPM, 230 Volts |
| Heating System | - Specially designed U shape S. S. tubular heating elements used on side walls to maintain uniform condition inside the chamber. - Heaters terminals taken out for wiring to avoid burning of heater terminals on high temperature. |
| Safety PID Controller | - Safety PID controller to cut-off the Heater supply in case of temperature overshoot with Alarm System. - Limit switch for interlocking connection between Motor & Door. i.e. when door is open; motor will be switch off. |
| Door | Single door fitted on heavy Patti hinges with Toggle clamp lock and Ceramic gasket. |
| Trays | Specially designed Heavy stainless steel rod trays ensure uniform temperature distribution. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside 14 SWG S.S. 304 with Heat resistance silver paint. Outer 16 SWG Mild steel (CRCA) sheet with hammer tone paint. |
| Temperature Range | 1000C to 4500C FRONT DOOR : 95 to 1000 C (Approx.) on 4500C Set Point. BOTH SIDE : 70 to 800 C (Approx.) |
| Temperature Accuracy | ± 0.50 C |
| Temperature Uniformity | + 100 C |
| Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |

| Models | Cap. Liters | Internal size H x W x D in cm | Trays | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI – 128 A HTF | 43 | 35 x 35 x 35 | 2 | 1,23,500 | 4,700 |
| TI – 128 B HTF | 91 | 45 x 45 x 45 | 2 | 1,50,200 | 4,700 |
| TI – 128 C HTF | 122 | 60 x 45 x 45 | 3 | 1,70,500 | 5,800 |
| TI – 128 D HTF | 216 | 60 x 60 x 60 | 3 | 1,98,500 | 5,800 |

NOTE: WHEN HIGH TEMPERATURE OVEN IS USE ABOVE 3000C THEN BLACK SPOTS OR STAINS GENERATED ON INNER CHAMBER, TO AVOID THE SAME WE APPLY HEAT RESISTANCE SILVER PAINT ON INNER CHAMBER.

Requirements for Installation :-

1.    PC requirement (if Software ordered) : 4 GB RAM, Windows 7 or 8,  CD ROM, 2 Nos. USB Port.

2.    Place the chamber where ventilation is good.

3.    The chamber must be operated in an ambient of 220 C ~ 300C.

4.    Leave minimum 50mm space from the sides and leave 300mm space from the backside.

5.    Place the chamber horizontally leveled.

6.   Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a

voltage stabilizer (KVA).

7.    Proper grounding of the chamber to get satisfactory operation of chamber.

8.    Do not test any product in the chamber, which is capable of generating combustible mixtures.

VACUUM OVEN  :-

| Make | “TEMPO” |  |
| --- | --- | --- |
| HSN Code | 8514 |  |
| Certification | CE Certification |  |
| Construction | Vacuum Oven is made of Stainless still (S.S.304 / S.S 316) inner and thick gauge Mild steel (CRCA) / S.S. 304 sheet duly epoxy powder coated outer body. Temperature achieved by heaters mounted outside the chamber to consume working chamber. Vacuum Valve & Gas inert line provided on one side of the unit. |  |
| Insulation | (Circular) – 3” High grade glass wool insulation to minimize heat loss. (Rectangular) – 4” High grade glass wool insulation to minimize heat loss. |  |
| Heating System | Heating element made from the best quality kanthal DSD wire is used to develop unique temperature inside the chamber. |  |
| Safety | Safety cum stand by thermostat to cut off the Heater supply in case of overshoot of temperature. |  |
| Door | Toughened glass window in door to allow viewing of samples inside the chamber. Secure door latch and continuous silicon gasket on vacuum chamber provide a tight seal & maintain. |  |
| Trays | Stainless steel shelf systems improve heat conduction to samples & easily removed for cleaning. |  |
| Supply | 230V AC, Single Phase, 50Hz |  |
| SPECIFICATIONS : |  |  |
| Standard Model |  | Inside S.S. 304 & Outer Mild steel (CRCA) sheet with powder coated. |
| Inner & Outer SS 304 |  | Inside S.S. 304 & Outer S.S. 304 |
| G.M.P. Model |  | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range |  | 500C to 2000C |
| Temperature Accuracy |  | ± 0.2 0 C |
| Control System |  | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |
| Vacuum Capacity |  | Can withstand a vacuum from 200 to 760 mm Hg of mercury relevant to height from sea level. |
| Vacuum PUMP |  | NOTE : Vacuum PUMP will be in your scope to be arranged during Installation |

| Models | Cap. Ltr. | Internal size H x W x D cm | Trays | Leak Rate | Standard Model (Amt. in Rs.) | Inner& Outer SS 304 (Amt. in Rs.) | G.M.P. Model (Amt. inRs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- | --- | --- |
| TI-135 A Circular | 15 | Depth 30 x 25 Cm Dia. With Glass window of 7”Dia | 2 | 25 mm / 24 hrs. | 48,200 | 72,800 | 79,600 | 2,000 |
| TI-135 B Circular | 28 | Depth 40 x 30 Cm Dia. With Glass window of 7”Dia | 2 | 25 mm / 24 hrs. | 55,600 | 91,800 | 97,600 | 2,200 |
| TI-136 GWA | 43 | 35 x 35 x 35 With Glass window | 2 | 50 mm / 24 hrs. | 73,300 | 93,700 | 99,600 | 3,500 |
| TI-136 GWB | 91 | 45 x 45 x 45 With Glass window | 2 | 50 mm / 24 hrs. | 1,01,200 | 1,28,600 | 1,39,400 | 4,500 |
| TI-136 GWC | 122 | 60 x 45 x 45 With Glass window | 3 | 50 mm / 24 hrs. | 1,08,700 | 1,46,800 | 1,56,500 | 6,000 |

NOTE : Initially on low temperature set point from 600C to 900 C, temperature over shoot & undershoot will be Approx 100C every time for fresh cycle start. After that temperature set point will be achieved.

| Documentation | : | We will provide IQ, OQ Documents. |
| --- | --- | --- |
| VALIDATION | : | We do not provide VALIDATION for Vacuum Oven, as such same does not have forced air circulation hence UNIFORMITY is not possible in Vacuum Oven. |

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Safety digital temperature controller with accessories | 6,000 |
| 4 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 5 | Digital Vacuum Indicator | 20,000 |

Requirement for Installation:

- Vacuum Pump required –Double stage with 250 LPM, 1 HP, 1440 RPM, Direct drive

NOTE: Vacuum PUMP will be in your scope to be arranged during Installation (in Customer Scope).

- PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

- Place the chamber where ventilation is good.  The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 50mm space from the sides and leave 300mm space from the backside.

- Place the chamber r horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

TRAY DRYING OVEN   (12 TRAYS) :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | - Double Wall Construction with heavy angle. - All sides of the Oven will be insulated with 7.5 cm thick Glass wool insulation to minimize skin temperature. - Design of panels will be such that inner & outer wall contact to be minimize heat losses through conduction. |
| Hot air circulation | 1/12 HP single phase motors (Two nos.) with fan to develop cross air flow circulation for better drying effect & uniformity of temperature. |
| Control System | - Microprocessor based PID Temperature controller. - Digital Timer: 0 to 999 (Hr. / Min. / sec.) with auto & manual position will be provided with the unit. - Thermostat - Heater High / Low switch. - MCB for mains |
| Heating System | S. S. Tubular air heaters on both sides. |
| Exhaust | Adjustable exhaust at top. |
| Double Door | Standard Model: 2 Nos. of Doors provided for the Tray Dryer with Asbestos Gasket. Inner & Outer SS 304 : 2 Nos. of Doors provided for the Tray Dryer with Silicon Gasket. GMP Model: 2 Nos. of Doors provided for the Tray Dryer with Silicon Gasket. |
| Trays | Stainless Steel trays (Plain or S.S. Rod tray) Note : Load caring capacity 5 kg (each tray) |
| Safety | - Safety thermostat to cut off the Heater supply in case of overshoot of temperature. - Limit switch will be provided on front door, which will cut-off supply as soon as door open. - MCB for mains |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S 304 & Outer Mild steel sheet with Hammer ton paint. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304 |
| GMP Mode | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range | 500C to 2000C |
| Temperature Accuracy | ± 10 C |
| Control System | Microprocessor based auto tuned PID controller cum indicator with dual display of set value & process value. |

| Inner Size H x W x D in cm | Outer Size H x W x D in cm | Cap. Ltr. | Tray Size H x W x D in cm | No. of Trays | Rating (KW) |
| --- | --- | --- | --- | --- | --- |
| 84 x 84 x 43 | 121 x 65 x 78 | 303 | 3 x 81 x 40 | 12 | 6 KW |

| Models | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | GMP Mode (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- |
| TI-12 Tray | 1,76,400 | 2,24,700 | 2,67,200 | 10,000 |

STANDARD MODEL WITH SILICON GASKET WILL BE CHARGE EXTRA.

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Safety digital temperature controller with accessories | 6,000 |

OUR OFFER EXCLUDES:

- All civil & structural work for installing the unit.

- Electrical work & cables up to our terminal point.

- Exhaust ducting.

- Any other items not specifically mentioned in our offer.

Requirement for Installation:

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1feet space from the sides and leave 2feet space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (440V +5%, 50Hz, 3 Phase).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

YOU WILL PROVIDE THE FOLLOWING :-

- Level  Floor

- Shifting of equipment at site

- Electrical  wiring  up  to  the  equipment  through  suitable  isolation  switch  and  to various  motors.

TRAY DRYING OVEN   (24 TRAYS) :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | - Double Wall Construction with heavy angle. - All sides of the Oven will be insulated with 7.5 cm thick Glass wool insulation to minimize skin temperature. - Design of panels will be such that inner & outer wall contact to be minimize heat losses through conduction - 1 no. Trolley track will be provided at bottom. |
| Hot air circulation | 1/6 HP single phase motors (Two nos.) with fan to develop cross air flow circulation for better drying effect & uniformity of temperature. |
| Control System | - Microprocessor based PID Temperature controller. - Digital Timer: 0 to 999 (Hr. / Min. / sec.) with auto & manual position will be provided with the unit. - Thermostat - MCB for mains - MCB for Booster Heater |
| Heating System | S. S. Tubular air heaters on both sides. |
| Exhaust | Adjustable exhaust at top. |
| Double Door | Standard Model: 2 Nos. of Doors provided for the Tray Dryer with Asbestos Gasket. Inner & Outer SS 304 : 2 Nos. of Doors provided for the Tray Dryer with Silicon Gasket. GMP Model: 2 Nos. of Doors provided for the Tray Dryer with Silicon Gasket. |
| Trolley & Trays | - Stainless Steel trays. (Plain Or S.S rod tray) - 1 no. trolley are provided with Tray dryer. - Total 24 trays. |
| Safety | - Safety thermostat to cut off the Heater supply in case of overshoot of temperature. - Limit switch will be provided on front door, which will cut-off supply as soon as doors open. - MCB for mains |
| Supply | 415V AC, Three Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S 304 & Outer Mild steel sheet with Hammer ton paint. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304. |
| GMP Model | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range | 500C to 2000C |
| Temperature Accuracy | ± 10 C |
| Control System | Microprocessor based auto tuned PID controller cum indicator with dual display of set value & process value. |

| Internal size H x W x D in cm | Cap. Ltr | No. of Trays | No. of Trolley | Tray Size H x W x D in cm | Rating (KW) |
| --- | --- | --- | --- | --- | --- |
| 96 x 100 x 92 | 883 | 24 | 1 | 3 x 40 x 81 | Total - 9 KW Main Heater - 6 KW Booster Heater - 3 KW |

| Models | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | GMP Model (Amt. in Rs.) | Packing Changes |
| --- | --- | --- | --- | --- |
| TI-24 Tray | 4,70,400 | 5,46,000 | 6,75,200 | 20,000 |

STANDARD MODEL WITH SILICON GASKET WILL BE CHARGE EXTRA.

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Safety digital temperature controller with accessories | 6,000 |

Our Offer Excludes:

- All civil & structural work for installing the unit.

- Electrical work & cables up to our terminal point.

- Exhaust ducting.

- Any other items not specifically mentioned in our offer.

Requirement for Installation:

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1feet space from the sides and leave 2feet space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (440V +5%, 50Hz, 3 Phase).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

YOU WILL PROVIDE THE FOLLOWING :-

- Level  Floor

- Shifting of equipment at site

- Electrical  wiring  up  to  the  equipment  through  suitable  isolation  switch  and  to various  motors

TRAY DRYING OVEN  (48 TRAYS) :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | - Double Wall Construction with heavy angle. - All sides of the Oven will be insulated with 7.5 cm thick Glass wool insulation to minimize skin temperature. - Design of panels will be such that inner & outer wall contact to be minimize heat losses through conduction. - 2 nos. Trolley track will be provided at bottom. |
| Hot air circulation | 1 HP single phase motors (two nos.) with fan to develop cross air flow circulation for better drying effect & uniformity of temperature. |
| Control System | - Microprocessor based PID Temperature controller. - Digital Timer: 0 to 999 (Hr. / Min. / sec.) with auto & manual position will be provided with the unit. - Thermostat - MCB for Mains - MCB for Booster Heater |
| Heating System | S. S. Tubular air heaters on both sides. |
| Exhaust | Adjustable exhaust at top. |
| Double Door | Standard Model : 2 Nos. of Doors provided for the Tray Dryer with Asbestos Gasket. Inner & Outer SS 304 : 2 Nos. of Doors provided for the Tray Dryer with Silicon Gasket. GMP Model : 2 Nos. of Doors provided for the Tray Dryer with Silicon Gasket. |
| Trolley & Trays | - Stainless Steel trays. (Plain Or S.S rod tray) - 2 nos. trolley are provided with Tray dryer. Each trolley will carry 24 nos. trays. - Total 48 trays. |
| Safety | - Safety thermostat to cut off the Heater supply in case of overshoot of temperature. - Limit switch will be provided on front door, which will cut-off supply as soon as doors open. - MCB for mains |
| Supply | 415V AC, Three Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S 304 & Outer Mild steel sheet with Hammer ton paint. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304. |
| GMP Model | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range | 50o C to 200o C |
| Temperature Accuracy | ± 10 C |
| Control System | Microprocessor based auto tuned PID controller cum indicator with dual display of set value & process value. |

| Internal size H x W x D in cm | Cap Ltr. | No. of Trays | No. of Trolley | Tray Size H x W x D in cm | Rating (KW) |
| --- | --- | --- | --- | --- | --- |
| 175 x 102 x 92 | 1642 | 48 | 2 | 3 x 40 x 81 | Total - 18 KW Main Heater - 12 KW Booster Heater - 6 KW |

| Models | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | G.M.P. Model (Amt. in Rs.) | Freight Changes |
| --- | --- | --- | --- | --- |
| TI-48 Tray | 7,37,100 | 8,90,400 | 11,13,000 | Door Delivery Special Vehicle |

STANDARD MODEL WITH SILICON GASKET WILL BE CHARGE EXTRA.

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Safety digital temperature controller with accessories | 6,000 |

Our Offer Excludes:

- All civil & structural work for installing the unit.

- Electrical work & cables up to our terminal point.

- Exhaust ducting.

- Any other items not specifically mentioned in our offer.

Requirement for Installation:

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1feet space from the sides and leave 2feet space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (440V +5%, 50Hz, 3 Phase).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

YOU WILL PROVIDE THE FOLLOWING :-

- Level  Floor

- Shifting of equipment at site

- Electrical  wiring  up  to  the  equipment  through  suitable  isolation  switch  and  to various  motors

DRUM HEATING OVEN :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Construction | The base heater for drums in the vertical position and is suitable for sue with type TI-129 (B & C) drum heaters. The heating power is 1 kw. At 240C AC single phase |
| SPECIFICATION: | Tempo drum heaters provide an easy, efficient, clean and economical way for the handling of viscous of low melting point soliditied materials in drums. Such materials include Asphalt, Bitumen, Fats, Gelatin, Glucose, Grease, Lubricating oils, Paints, Petroleum Jelly, Phenols, Plastics, Sodium Tar, Vanaspati, Oil, Varnishes & Waxes etc. Tempo drum heaters are mobile ovens heated by “Tempanel” flexible surface electric heating elements operating at black heat. They are mode in rugged designs and stand normal usage in chemical industries. They can be used for heating the contents of drums kept in the vertical or horizontal positions. The contents will be heated at a rate depending on the size of the drum and its contents. The contents may either be melted completely and removed as liquid or they may be melted only at the surface touching the drum walls and removed as a slid block by lifting the drum. The heating in tempo Drum Heaters in gentle and uniform and there is no chance of local overheating. Automatic temperature control is provided as standard equipment, except in the base and top heaters in which it is not required. We can also make special HEATING JACKETS for most of your heating applications in your laboratory, pilot plant or process line e.g. Jackets for heating laboratory glassware, reaction vessels, kettles, special size drums, metal pipelines etc. |
| Base Heater | For TI- 129A – This is a base heater for drums in the vertical position and is suitable for use with type TI-129 (B & c) Drum Heaters. The heating power is 1 KW at 240 V AC. Single Phase. Base : Inner Dia. 2 ft. |
| Vertical Drum Heater | TI-129B / C is a vertical drum heater in two sections hinged to each other, forming an cylindrical heater when closed. It is fitted with castor wheels. The standard model can take a standard size drum of 45 gallons capacity. |
| Supply | For TI 129 B 230 V Single Phase AC |
|  | For TI 129 C 440 V Three Phase AC |
| Control System (For 129 B & C) | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |

| Models | Inner Dia & height | Rating KW | Maximum temp. | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-129 A (Base Heater) | 2ft. Dia. | 1 | Suitable for use with TI-129 B / TI-129 C | 18,900 | 750 |
| TI-129 B (Drum Heater) | 2 ft & height 3 ft. | 6 | 2500C | 56,800 | 3,700 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

LEAK TEST APPRATUS (SEMI AUTOMATIC):-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8479 |
| Certification | CE Certification |
| Construction | Leak Test Apparatus is made of Stainless still (S.S.304) sheet. Vacuum archived achieved by pump mounted inside the chamber. Vacuum Valve & Timer provided on front side of chamber. Vacuum line provided on back side of the unit. |
| Vacuum System | Fast vacuum created by pump from the best quality of Diaphragm vacuum pump. |
| Desiccator | Desiccator is made up of polycarbonate Material which allow viewing of samples inside the chamber |
| Timer | Electrical Timer with digital display and settable time. |
| Buzzer and lamp | After completion of cycle Buzzer will give alarm with Indication lamp |
| Vacuum isolating valve | Isolating valve used for hold the vacuum at exact level. |
| Vacuum Display | S.S. Vacuum Gauge |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Construction | S.S. 304 dull smooth buff finish. |
| Timer | Digital timer with settable time. |
| Vacuum Capacity | Can withstand a vacuum of - 550 mm of mercury. |
| Vacuum PUMP | Oil free Diaphragm pump |

| Models | Internal size | Inner & Outer SS 304 (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- |
| TI-140 A | Desiccator 6”Dia X 9” Ht | 48,800 | 1,500 |
| TI-140 B | Desiccator 12”Dia X 11” Ht | 55,600 | 1,500 |

| Documentation | : | We will provide IQ, OQ & PQ Documents. Calibration certificate & SOP |
| --- | --- | --- |

LEAK TEST APPRATUS (FULLY AUTOMATIC) :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8479 |
| Certification | CE Certification |
| Construction | Leak Test Apparatus is made of Stainless still (S.S.304) sheet. Vacuum archived achieved by pump mounted inside the chamber. Vacuum line provided on back side of the unit. |
| Vacuum System | Fast vacuum created by pump from the best quality of Diaphragm vacuum pump. |
| Desiccator | Desiccator is made up of polycarbonate Material which allow viewing of samples inside the chamber |
| Controller | - LCD display with programmable set Vacuum & Hold Time as per requirement. - Soft touch membrane keypad |
| Buzzer and lamp | After completion of cycle Buzzer will give alarm with Indication lamp |
| Solenoid valve | Solenoid valve used for hold the vacuum at exact level. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Construction | S.S. 304 dull smooth buff finish. |
| Vacuum Capacity | Can withstand a vacuum of - 600 mm of mercury. |
| Vacuum PUMP | Oil free Diaphragm pump |
| Features | - Large size numerical LED display. - Portable & Low power consumption. - Automatically release of Vacuum from the chamber at the end of cycle. - Continuous display of time & vacuum level. - Vacuums & time can be changed. - Direct printer (EPSON LX300+) Interface. |

| Models | Size | Inner & Outer SS 304 (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- |
| TI-141 A | Desiccator 6”Dia x 9” Ht | 78,000 | 1,500 |
| TI-141 B | Desiccator 12”Dia x 11” Ht | 84,800 | 1,500 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

| Documentation | : | We will provide IQ, OQ & PQ Documents. Calibration, SOP |
| --- | --- | --- |

LAB CLEAN POSITIVE PRESSURE SYSTEM :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8421 |
| Certification | CE Certification. |
| Air recirculation | Blower at bottom which ensure uniform filtered airflow. |
| Blower motor | 1440 RPM, 230 Volts |

| SPECIFICATION : |  |
| --- | --- |
| Construction | Single wall construction and available is S.S. & G.I. Model. |
| Three Stage Filtration | PRE FILTER-Synthetic PU Washable 5 Micron |
|  | HEPA FILTER- Efficiency 99.97% @ 0.3 microns |
|  | VOC FILTER-Unique Blend of Activated Carbon and Alumina |
| CFM capacity | 350 CFM |
| Mounting | Floor Standing with wheels, Front Lockable. |
| Power consumption | 100 Watts Maximum |
| Filter replacement | Unique Automatic Indicator for Changing Filter. |
| Electricity supply | 220 / 240 Volts 50 / 60 Hz. |
| Standard Model | G.I. Powder Coated |
| GMP Model | Stainless Steel Model - S.S. 304 |
| Noise level | Less than 65 DB |

| Model | Unit Capacity | Area Covered | Standard Model (Amt. in Rs.) | GMP Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-350 LC | 350CFM | 350 sq. feet | 66,200 | 92,600 | 2,500 |

Documentation : We will provide   HEPA Filter Test Certificate, PRE Filter Test Certificate, VOC Filter Test Certificate, Motor – Blower Test Certificate & Material of Construction Certificate for SS 304.

Warranty : 1 Year warranty on Blower and Timer.

Requirement for Installation

- Leave minimum 50mm space from the all sides.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V 5%, 50Hz, 1 Phase) Its recommended to use a voltage stabilizer (KVA)

- Proper grounding of the chamber to get satisfactory operation of chamber.

BACTERIOLOGICAL INCUBATOR (Natural Convection):-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | Double walled construction |
| Insulation | 3” Thick PUF Insulation ensures stable temperature & reduced energy consumption. |
| Heating System | Heating elements made from the best quality Kanthal DSD wire are used on refractory supports. |
| Observation Door | Full view Acrylic door / Glass Door with gasket to view samples inside the chamber. |
| Trays | Specially designed stainless steel rod trays ensure uniform temperature distribution. |
| Safety | Additional Safety cum stand by thermostat to cut of heater supply in case temperature of overshoot. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 mirror finish & Outer Mild steel (CRCA) sheet with powder coated. |
| Temperature Range | 50C above ambient to 600C |
| Temperature Accuracy | ± 0.5 0C |
| Control System | Imported microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |

| Models | Capacity Liters | Internal size H x W x D cm | Trays | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-90 A | 43 | 35 x 35 x 35 | 2 | 34,500 | 2,000 |
| TI-90 B | 91 | 45 x 45 x 45 | 2 | 44,100 | 3,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

Requirement for Installation:

- Place the chamber where ventilation is good.  The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 300 mm space from the sides and leave 300mm space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

BACTERIOLOGICAL INCUBATOR with forced air recirculation  :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | Double walled construction with backside triple. |
| Insulation | 3” Thick PUF Insulation ensures stable temperature & reduced energy consumption. |
| Air Circulation | Motorized Blower on back side of the chamber develops unique air flow system which ensures maximum uniform temperature distribution inside the chamber. Unique air flow assures quick recovery after door openings. |
| Blower motor | 1/12 HP, TEFC , F-class insulation, Single phase, 1440 RPM, 230 Volts |
| Heating System | High quality S.S. Tubular Heaters are used for better heating conditions. |
| Observation Door | Full view Acrylic door / Glass Door with gasket to view samples inside the chamber. |
| Trays | Specially designed stainless steel rod trays ensure uniform temperature distribution. |
| Safety | Additional Safety cum stand by thermostat to cut of heater supply in case temperature of overshoot. |
| Features | - Fiber/PU Wheel with front lockable for easy movement. (available for 100 Ltr. Capacity onwards) - MCB for electrical safety - Adjustable tray height arrangements. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 & Outer Mild steel (CRCA) sheet with powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 & Outer S.S. 304 |
| G.M.P. Model | Inside S.S. 316 & Outer S.S. 304 |
| Temperature Range | 50C above ambient to 600C SPL. Note: While operating the machine on set point of 350C above & up to 600C; Room Temperature must have to be maintain between 250C ~ 270C. |
| Temperature Accuracy | ± 0.2 0C |
| Temperature Uniformity | ± 1 0C |
| Control System | Imported microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |

| Models | Capacity Liters | Internal size H x W x D cm | Max. No. of Trays | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | GMP Model (Amt. in Rs.) | Packing Charge |
| --- | --- | --- | --- | --- | --- | --- | --- |
| Table Top models |  |  |  |  |  |  |  |
| TI-90 A | 43 | 35 x 35 x 35 | 2 | 51,200 | 66,800 | 69,000 | 2,500 |
| TI-90 B | 91 | 45 x 45 x 45 | 2 | 61,800 | 84,900 | 87,100 | 3,300 |
| Floor models with Wheels |  |  |  |  |  |  |  |
| TI-90 C | 100 | 50 x 50 x 40 | 2 | 72,400 | 89,900 | 96,000 | 3,700 |
| TI-90 D | 170 | 70 x 60 x 40 | 2 | 82,100 | 1,06,600 | 1,13,900 | 3,700 |
| TI-90 E | 200 | 60 x 60 x 60 | 2 | 88,900 | 1,18,100 | 1,26,000 | 3,700 |
| TI-90 F | 285 | 80 x 60 x 60 | 3 | 1,00,000 | 1,33,100 | 1,43,600 | 5,000 |
| TI-90 G | 325 | 90 x 60 x 60 | 4 | 1,07,700 | 1,44,400 | 1,56,200 | 5,000 |
| TI-90 H | 400 | 110 x 60 x 60 | 4 | 1,16,200 | 1,68,000 | 1,82,000 | 7,500 |
| TI-90 I | 600 | 125 x 70 x 70 | 4 | 1,50,800 | 2,00,600 | 2,29,400 | 9,000 |
| TI-90 J | 800 | 125 x 80 x 80 | 5 | 1,69,400 | 2,23,000 | 2,51,600 | 11,000 |
| TI-90 K | 1000 | 157 x 80 x 80 | 5 | 1,83,000 | 2,45,400 | 2,80,500 | 15,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

Note	               :  Two blowers for 800 ltrs to 1000 ltrs. Capacity chambers.

Validation & Documentation :  2 hrs. Loaded OR unloaded condition on one cycle. (4 or 8 temp. probes).

Along with IQ, OQ & PQ Documents.

Cost of Validation	               :  EXTRA Rs.  15,000/- + GST @ 18% Extra.

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Scanner : Scanner 4 or 8 point temperature scanner complete with sensors. | 33,000 |
| 4 | Safety digital temperature controller with accessories | 6,000 |
| 5 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 6 | Stainless steel rod trays : 100 ltr. to 400 ltr. | 3,000 |
| 7 | Stainless steel rod trays : 600 ltr. to 1000 ltrs. | 4,500 |

Requirements for Installation :-

- PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 300 mm space from the sides and leave 300mm space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

MUFFLE FURNACE   930 0C  :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Construction | Horizontal rectangular muffle. |
| Heating System | Heating is affected by means of element made of best quality Nichrome wire wounded on the muffle. |
| Insulation | Heating element wounded with high grade cerewool blanket and the gaps are filled with glass wool / cerewool to avoid heat loss. |
| Safety | Energy regulator to control heating elements for extra heating and safety of equipment. |
| Door | Single door made of brick work and cementing fitted in with single hole at Center. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inner made of ceramic pot. Outer case made from thick gauge Mild Steel (CRCA) sheet. |
| G.M.P Model | Inner made of ceramic pot. Outer case made from thick gauge Stainless Steel (S.S. 304) sheet. |
| Temperature Range | Maximum working temperature up to 9300 C. NOTE : Especially in a GMP Model, Outer is made of Stainless Steel, it is not possible to achieve low skin temperature as compare to STD Model which have heat resistant thick paint applied on it. FRONT DOOR : 105 to 1200 C (approx.) on 9000C Set Point. BOTH SIDE : 80 to 950 C (approx.) |
| Temperature Accuracy | ± 30C |
| Temperature Uniformity | ± 100C |
| Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |

| Models | Cap. Liters | Internal size H x W x D in cm | Rating KW | Standard Model (Amt. in Rs.) | G.M.P. Model (Amt. in Rs.) | Packing Charge |
| --- | --- | --- | --- | --- | --- | --- |
| TI-58 A | 2.3 | 10 x 10 x 23 | 1.6 | 34,700 | 48,300 | 2,500 |
| TI-58 B | 6.75 | 15 x 15 x 30 | 3.3 | 48,300 | 66,200 | 3,500 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Safety digital temperature controller with accessories | 6,000 |
| 4 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 5 | Profile controller / Programmable controller - 8 ramp & 8 sack period | 16,000 |

Requirements for Installation :-

- PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

- Place the chamber where ventilation is good.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

MUFFLE FURNACE HIGH TEMPERATURE (With Thyrister Drive) :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Construction | Horizontal rectangular ceramic muffle |
| Heating System | Heating is affected by means of element made of best quality. Kanthal A-1 wire wounded on the muffle. |
| Insulation | Heating element wounded with high grade cerewool blanket and the gaps are filled with cerewool to avoid heat loss. |
| Safety Controller | Safety digital temperature controller with alarm |
| Door | Single door made of brick work and cementing fitted in with single hole at Center. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inner made of ceramic pot. Outer case made from thick gauge Mild Steel (CRCA) sheet |
| G.M.P Model | Inner made of ceramic pot. Outer case made from thick gauge Stainless Steel (S.S. 304) sheet. |
| Temperature Range | Max. working temp. 11500C for Model No. TI-58 HTA, 58- HTB & 58-HTC Max. working temp. 10000C for Model No. TI-58 HTD NOTE : Especially in a GMP Model, Outer is made of Stainless Steel, it is not possible to achieve low skin temperature as compare to STD Model which have heat resistant thick paint applied on it. FRONT DOOR : 105 to 1200 C (approx.) on 10000C Set Point. BOTH SIDE : 80 to 950 C (approx.) |
| Temperature Accuracy | ± 30C |
| Temperature Uniformity | ± 100C |
| Control System | Imported microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |
| Thyrister Drive | An additional Thyrister Circuit reduces energy consumption which results into decrease in electricity bills. |

| Models | Capacity Liters | Internal size H x W x D in cm | Rating KW | Standard Model (Amt. in Rs.) | GMP Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| MAXIMUM WORKING TEMPERATURE 11500C |  |  |  |  |  |  |
| TI-58HTA | 2.3 | 10 x 10 x 23 | 1.8 | 47,700 | 60,100 | 2,500 |
| TI-58HTB | 6.75 | 15 x 15 x 30 | 3.5 | 66,200 | 82,200 | 4,000 |
| TI-58HTC | 12 | 20 x 20 x 30 | 3.5 | 79,400 | 96,600 | 4,500 |
| MAXIMUM WORKING TEMPERATURE 10000C |  |  |  |  |  |  |
| TI-58HTD | 22.78 | 22.5 x 22.5 x 45 | 7 | 1,08,800 | 1,26,000 | 6,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

Validation & Documentation :  2 hrs. on one set point  (2 temperature probes) along with IQ, OQ & PQ Documents.

Cost of Validation	               :  EXTRA Rs.  15,000/- + GST @ 18% Extra.

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 4 | Profile controller / Programmable controller - 8 ramp & 8 sack period | 16,000 |

Requirements for Installation :-

- PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

- Place the chamber where ventilation is good.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

Advantages of TEMPO make Muffle Furnace with Thyrister drive PID Controller :-

Temperature Range: -

Maximum Working Temperature 11500C.

- Construction : Tempo is using Heavy guage metal fabrication to with stand high Temp. Also main body

where door of Furnace rest is manufactured in 3mm Thick Plate to with stand high temp & long life.

- We use imported CE mark PID Controller with Thyrister Drive controlling system which is a best safety

system for Heating Element of Furnace.

Thyrister Drive circuit Profile Controller enhances following benefits:-

- Reduces Energy consumption which results into decreasing electricity bills.

- Heating Elements safety increases against variation in electric supply.

- Thyrister safety CKT increase lifer of Heating Element.

- Also note that we are using Kanthal A1 imported coil as Heating Element

MUFFLE FURNACE HIGH TEMPERATURE with Shell :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Construction | Horizontal rectangular ceramic muffle. Control panel fitted on bottom side. |
| Outer Shell | Shell with air gap for low body temperature. Outer shell minimize outer skin Temperature of Muffle Furnace. |
| Heating System | Heating is affected by means of element made of best quality. Kanthal A-1 wire wounded on the muffle. |
| Insulation | Heating element wounded with high grade cerewool blanket and the gaps are filled with cerewool to avoid heat loss. |
| Safety Controller | Safety digital temperature controller with alarm |
| Door | Single door made of brick work and cementing fitted in with single hole at Center Door. |
| Supply | 230V AC, Single Phase, 50Hz also with False covershell. |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inner made of ceramic pot. Outer case made from thick gauge Mild Steel (CRCA) sheet with mild steel outer shell. |
| G.M.P Model | Inner made of ceramic pot. Outer shell made from thick gauge Stainless Steel (S.S. 304) sheet. |
| Temperature Range | Max. working temp. 11500C for Model No. TI-59 HTA, 59-HTB & 59-HTC Max. working temp. 10000C for Model No. TI-59 HTD NOTE : Especially in a GMP Model, Outer is made of Stainless Steel, it is not possible to achieve low skin temperature as compare to STD Model which have heat resistant thick paint applied on it. |
| Temperature Accuracy | ± 30C |
| Temperature Uniformity | ± 100C |
| Control System | Imported microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |
| Thyrister Drive | An additional Thyrister Circuit reduces energy consumption which results into decrease in electricity bills. |

| Models (S) or (G) | Cap. Liters | Internal size H x W x D in cm | Rating KW | Standard Model (Amt. in Rs.) | GMP Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| MAXIMUM WORKING TEMPERATURE 11500C |  |  |  |  |  |  |
| TI-59 HTA | 2.3 | 10 x 10 x 23 | 1.8 | 55,100 | 75,400 | 2,500 |
| TI-59 HTB | 6.75 | 15 x 15 x 30 | 3.5 | 75,000 | 98,700 | 4,000 |
| TI-59 HTC | 12 | 20 x 20 x 30 | 3.5 | 88,200 | 1,18,900 | 4,500 |
| MAXIMUM WORKING TEMPERATURE 10000C |  |  |  |  |  |  |
| TI-59 HTD | 22.78 | 22.5 x 22.5 x 45 | 7 | 1,19,100 | 1,51,200 | 6,000 |

NOTE : GMP Model is available with mild steel shell price on request.

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,000 |
| 4 | Profile controller / Programmable controller - 8 ramp & 8 sack period | 16,000 |

Requirements for Installation :-

- PC requirement (if Software ordered) : 8 GB RAM, Windows 8, 10 or 11,  CD ROM, 2 Nos. USB Port, Ethernet Port.

- Place the chamber where ventilation is good.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

Advantages of TEMPO make Muffle Furnace with Thyrister drive PID Controller :-

Temperature Range: -

Maximum Working Temperature 11500C.

- Construction : Tempo is using Heavy guage metal fabrication to with stand high Temp. Also main body

where door of Furnace rest is manufactured in 3mm Thick Plate to with stand high temp & long life.

- We use imported CE mark PID Controller with Thyrister Drive controlling system which is a best safety

system for Heating Element of Furnace.

Thyrister Drive circuit Profile Controller enhances following benefits:-

- Reduces Energy consumption which results into decreasing electricity bills.

- Heating Elements safety increases against variation in electric supply.

- Thyrister safety CKT increase lifer of Heating Element.

- Also note that we are using Kanthal A1 imported coil as Heating Element

ECONOMY MODEL  FURNACE - HIGH TEMPERATURE 1300 0C :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Maximum Temp. Range | 13000C |
| Continuous working Temp. | 12500C |
| Body Construction | MS fabricated with duly powder coated, dual body structure. |
| Door | Leak proof Swing aside type door with safety switch |
| Insulation | Low energy consumption vacuum formed fiber boards & refractory bricks, Dual wall insulation, Hot wall insulation will be 14250C HD grade refractory bricks |
| CONTROL SYSTEM |  |
| Heating Rate | Up to 100C/Min |
| Heating element | SiC, Heating from Top sides for better uniformity |
| Skin Temperature | Less than 600C at 12500C |
| Thermocouple (DUAL) | R type with alumina sheath (Main Controller 1 no. & Safety Controller 1 no.) |
| Temperature Controller | Digital Temperature Programmable PID controller. Single Programme with 30 segments |
| Safety Controller | Safety controller (With Separate Thermocouple as above) |
| Temperature Accuracy | +/- 30C |
| Temperature Uniformity | +/- 150C |
| Thyristor Unit | SCR unit will be Provided with soft starting driver |
| Control Switch | Rocker switch, To switch control on/off |
| MCB | MCB will be provided for over current and short circuit protection to the circuit. Also works as main incomer of the system |
| Ammeters | Digital Ammeters to indicate load current |
| Indication | LED indicators |
| FEATURES |  |
| - Rugged and durable construction. |  |
| - Outer casing is made up of thick mild steel with duly powder coated, gives long service life with extremely resistant. |  |
| - Multilayer insulation with vacuum formed ceramic fibre board and refractory bricks with low thermal mass. |  |
| - Parallel guided swing aside door design for providing ease of operation. |  |
| - Furnace design with powerful heating elements for providing fast heating rate. |  |
| - Fast heat up and cool down rates. |  |
| - Better Temperature Uniformity |  |
| - Temperature Controller with PID Controllers. |  |
| - Power control with Thyristor units for very precise temperature Control, wear-free and noiseless. |  |
| - High mechanical durability, short heating cycle and low power consumption. |  |
| - Furnace is designed with powerful heating elements, i.e, SiC on all sides of the chamber providing fast heating rate. |  |

| Models | Internal size H x W x D in cm | Outer Size H x W x D in cm | Rating KW, Supply & Heater | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-60 HTA | 10 x 10 x 23 |  | 1 Phase, 2 Wire, 230 Volts, 4.1 KW, | PRICE ON REQUEST | 6,500 |
| TI-60 HTB | 15 x 15 x 30 |  | 1 Phase, 2 Wire, 230 Volts, 4.6 KW, | PRICE ON REQUEST | 6,800 |
| TI-60 HTC | 20 x 20 x 30 |  | 1 Phase, 2 Wire, 230 Volts, 5.2 KW, | PRICE ON REQUEST | 7,500 |

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |

SPARE PARTS KEEP IN STOCK

|  |  | List Price (Amt. in Rs.) |
| --- | --- | --- |
| 1 | Heaters for TI-60 HTA (6 nos. Heater) | PRICE ON REQUEST |
| 2 | Heaters for TI-60 HTB (6 nos. Heater) | PRICE ON REQUEST |
| 3 | Heaters for TI-60 HTC (8 nos. Heater) | PRICE ON REQUEST |
| 4 | Thyristor for TI-60 HTA (Single Phase) | PRICE ON REQUEST |
| 5 | Thyristor for TI-60 HTB (Single Phase) | PRICE ON REQUEST |
| 6 | Thyristor for TI-60 HTC (Single Phase) | PRICE ON REQUEST |
| 7 | Digital Ammeter for TI-60 HTA (Single Phase) | PRICE ON REQUEST |
| 8 | Digital Ammeter for TI-60 HTB (Single Phase) | PRICE ON REQUEST |
| 9 | Digital Ammeter for TI-60 HTC (Single Phase) | PRICE ON REQUEST |
| 10 | R Type Thermocouple | PRICE ON REQUEST |
| 11 | Safety Controller (Power and safety contactors) | PRICE ON REQUEST |
| 12 | PID Controller (Main Controller) | PRICE ON REQUEST |
| 13 | Safety digital temperature controller (Thermocouple) | PRICE ON REQUEST |

- INSULATION CRACKING

The insulation material in this furnace is susceptible to surface cracking as a result of temperature cycling. This is

a normal occurrence and such cracking is not detrimental to the performance of the overall life of the furnace,

under normal operation.

- LOADING THE FURNACE

When heating large object, in particular poor conductors, avoid shielding the thermocouple from the heating

element. Also ensure that nothing is within 15 mm of the elements.

The thermocouple is intended to sense the temperature near the heating element.

However if a large object is placed in the chamber it may record the average

Temperature of the object and the elements, which can lead to over heating of the Elements.

Allow large objects to gain heat at a lower temperature and then reset the controller to a

Temperature close to the desired maximum.

Take care that nothing hits the elements when loading and unloading.

- OPENING THE DOOR

Take great care when loading or unloading the furnace chamber. See section 4.2 before removing a hot object

from the furnace; ensure a suitable surface is available on which to put it.

Do not open the door at high temperature. If possible, do the open at above 2000C. If it is necessary to load or

unload work at elevated temperatures, keep the door open for short a period as possible. The insulation cools

quickly and may crack through thermal shock.

The heating elements and the insulation are very susceptible to mechanical shock. At all times operate the door

gently to avoid mechanical shock.

Opening the door when the furnace is very hot can cause hot air to be sucked by the case cooling fans towards

the thermal cut out and activate it (see section 4.11). Avoid excessive door opening and shut it as soon as

possible.

WARRANTY:

- Heating Elements and Alumina Tube are not warranted.

- All the components of furnace are warranted for the duration of 12 Months against any manufacturing defect.

- The defected component will be replaced.

- The furnace should be used as per the instructions and training.

- Any damage occurring due to mishandling and wrong operational practice is not covered under warranty.

- In case of Quartz Tube Furnace, devitrification of quartz tube occurring from prolong use is not covered in warranty.

FURNACE - HIGH TEMPERATURE 1300 0C :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Maximum Temp. Range | 13000C |
| Continuous working Temp. | 12500C |
| Body Construction | Ms fabricated with duly powder coated, dual body structure. Furnace is mounted on a heavy-duty stand. |
| Door | Leak proof Swing aside type door with safety switch |
| Insulation | Low energy consumption vacuum formed fiber boards & refractory bricks, Dual wall insulation, Hot wall insulation will be 14250C HD grade refractory bricks |
| CONTROL SYSTEM |  |
| Heating Rate | Up to 100C/Min |
| Heating element | SiC, Heating from two sides for better uniformity |
| Skin Temperature | Less than 600C at 12500C |
| Thermocouple (DUAL) | R type with alumina sheath (Main Controller 1 no. & Safety Controller 1 no.) |
| Temperature Controller | Digital Temperature Programmable PID controller. Single Programme with 30 segments |
| Safety Controller | Safety controller (With Separate Thermocouple as above) |
| Temperature Accuracy | +/- 30C |
| Temperature Uniformity | +/- 150C |
| Thyristor Unit | SCR unit will be Provided with soft starting driver |
| Control Switch | Rocker switch, To switch control on/off |
| MCB | MCB will be provided for over current and short circuit protection to the circuit. Also works as main incomer of the system |
| Ammeters | Digital Ammeters to indicate load current |
| Indication | LED indicators |
| FEATURES |  |
| - Rugged and durable construction. |  |
| - Outer casing is made up of thick mild steel with duly powder coated, gives long service life with extremely resistant. |  |
| - Multilayer insulation with vacuum formed ceramic fibre board and refractory bricks with low thermal mass. |  |
| - Parallel guided swing aside door design for providing ease of operation. |  |
| - Furnace design with powerful heating elements for providing fast heating rate. |  |
| - Fast heat up and cool down rates. |  |
| - Better Temperature Uniformity |  |
| - Temperature Controller with PID Controllers. |  |
| - Power control with Thyristor units for very precise temperature Control, wear-free and noiseless. |  |
| - High mechanical durability, short heating cycle and low power consumption. |  |
| - Furnace is designed with powerful heating elements, i.e, SiC on all sides of the chamber providing fast heating rate. |  |

| Standard Model | Cap. Ltr. | Internal size H x W x D in Cm | Outer Size H x W x D in Cm | Rating KW, Supply & Heater | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| TI-61 HTA | 2.3 | 10 x 10 x 23 |  | 1 Phase, 230 Volts, 2 Wire, 4.3 KW, | PRICE ON REQUEST | 6,500 |
| TI-61 HTB | 6.75 | 15 x 15 x 30 |  | 3 Phase, 415 Volts, 4 Wire, 5.8 KW | PRICE ON REQUEST | 6,800 |
| TI-61 HTC | 12 | 20 x 20 x 30 |  | 3 Phase, 415 Volts, 4 Wire, 6.3 KW, | PRICE ON REQUEST | 7,500 |
| TI-61 HTD | 22.78 | 22.5 x 22.5 x 45 |  | 3 Phase, 415 Volts, 4 Wire 8 KW, | PRICE ON REQUEST | 9,000 |

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |

SPARE PARTS KEEP IN STOCK

|  |  | List Price Amt. in Rs. |
| --- | --- | --- |
| 1 | SiC Heating Elements for TI-61 HTA (Heaters Qty. 8 nos.) | PRICE ON REQUEST |
| 2 | SiC Heating Elements for TI-61 HTB (Heaters Qty. 12 nos.) | PRICE ON REQUEST |
| 3 | SiC Heating Elements for TI-61 HTC (Heaters Qty. 12 nos.) | PRICE ON REQUEST |
| 4 | SiC Heating Elements for TI-61 HTD (Heaters Qty. 18 nos.) | PRICE ON REQUEST |
| 5 | Thyristor for TI-61 HTA (Single Phase) | PRICE ON REQUEST |
| 6 | Thyristor for TI-61 HTB (Three Phase) | PRICE ON REQUEST |
| 7 | Thyristor for TI-61 HTC (Three Phase) | PRICE ON REQUEST |
| 8 | Thyristor for TI-61 HTD (Three Phase) | PRICE ON REQUEST |
| 9 | Digital Ammeter for TI-61 HTA (Single Phase) | PRICE ON REQUEST |
| 10 | Digital Ammeter for TI-61 HTB (Three Phase) | PRICE ON REQUEST |
| 11 | Digital Ammeter for TI-61 HTC (Three Phase) | PRICE ON REQUEST |
| 12 | Digital Ammeter for TI-61 HTD (Three Phase) | PRICE ON REQUEST |
| 13 | R Type Thermocouple (20 cm) | PRICE ON REQUEST |
| 14 | PID Controller (Main Controller) | PRICE ON REQUEST |
| 15 | Safety Controller (Power and safety contactors) | PRICE ON REQUEST |
| 16 | Safety digital temperature controller with Thermocouple | PRICE ON REQUEST |

- INSULATION CRACKING

The insulation material in this furnace is susceptible to surface cracking as a result of temperature cycling. This is

a normal occurrence and such cracking is not detrimental to the performance of the overall life of the furnace,

under normal operation.

- LOADING THE FURNACE

When heating large object, in particular poor conductors, avoid shielding the thermocouple from the heating

element. Also ensure that nothing is within 15 mm of the elements.

The thermocouple is intended to sense the temperature near the heating element.

However if a large object is placed in the chamber it may record the average

Temperature of the object and the elements, which can lead to over heating of the Elements.

Allow large objects to gain heat at a lower temperature and then reset the controller to a

Temperature close to the desired maximum.

Take care that nothing hits the elements when loading and unloading.

- OPENING THE DOOR

Take great care when loading or unloading the furnace chamber. See section 4.2 before removing a hot object

from the furnace; ensure a suitable surface is available on which to put it.

Do not open the door at high temperature. If possible, do the open at above 2000C. If it is necessary to load or

unload work at elevated temperatures, keep the door open for short a period as possible. The insulation cools

quickly and may crack through thermal shock.

The heating elements and the insulation are very susceptible to mechanical shock. At all times operate the door

gently to avoid mechanical shock.

Opening the door when the furnace is very hot can cause hot air to be sucked by the case cooling fans towards

the thermal cut out and activate it (see section 4.11). Avoid excessive door opening and shut it as soon as

possible.

WARRANTY:

- Heating Elements and Alumina Tube are not warranted.

- All the components of furnace are warranted for the duration of 12 Months against any manufacturing defect.

- The defected component will be replaced.

- The furnace should be used as per the instructions and training.

- Any damage occurring due to mishandling and wrong operational practice is not covered under warranty.

- In case of Quartz Tube Furnace, devitrification of quartz tube occurring from prolong use is not covered in Warranty.

FURNACE - HIGH TEMPERATURE 1450 0C :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Maximum Temp. Range | 14500C |
| Continuous working Temp. | 14000C |
| Body Construction | Ms fabricated with duly powder coated, dual body structure |
| Door | Leak proof Swing aside type door |
| Insulation | Vacuum formed fiber boards, triple wall insulation, Hot wall insulation will be 16000C HD grade |
| CONTROL SYSTEM |  |
| Heating Rate | Up to 100C/Min |
| Heating element | SiC, Heating from two sides for better uniformity |
| Skin Temperature | Less than 600C at 14000C |
| Thermocouple (DUAL) | R type with alumina sheath (Main Controller 1 no. & Safety Controller 1 no.) |
| Temperature Controller | Digital Temperature Programmable PID controller. Single Programme with 30 segments |
| Safety Controller | Safety controller (With Separate Thermocouple as above) |
| Temperature Accuracy | +/- 30C |
| Temperature Uniformity | +/- 150C |
| Thyristor Unit | SCR will be Provided with soft starting driver |
| Control Switch | Rocker switch, to switch control on/off |
| MCB | MCB will be provided for over current and short circuit protection to the circuit. Also works as main incomer of the system |
| Ammeters | Digital Ammeters to indicate load current |
| Indication | LED indicators |
| FEATURES |  |
| - Rugged and durable construction. |  |
| - Outer casing is made up of thick mild steel with duly powder coated, gives long service life with extremely resistant. |  |
| - Multilayer insulation with vacuum formed ceramic fibre board and refractory bricks with low thermal mass. |  |
| - Parallel guided swing aside door design for providing ease of operation. |  |
| - Furnace design with powerful heating elements for providing fast heating rate. |  |
| - Fast heat up and cool down rates. |  |
| - Better Temperature Uniformity |  |
| - Temperature Controller with PID Controllers. |  |
| - Power control with Thyristor units for very precise temperature Control, wear-free and noiseless. |  |
| - High mechanical durability, short heating cycle and low power consumption. |  |
| - Furnace is designed with powerful heating elements, i.e, SiC on all sides of the chamber providing fast heating rate. |  |

| Standard Model | Internal size H x W x D in cm | Outer size H x W x D in cm | Rating KW, Supply & Heater | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-62 HTA | 10 x 10 x 23 |  | 1 Phase, 230 Volts 2 Wire, 4.5 KW, | PRICE ON REQUEST | 7,500 |
| TI-62 HTB (SPL) | 15 x 15 x 20 |  | 1 Phase, 230 Volts, 2 Wire 4.2 KW, | PRICE ON REQUEST | 8,000 |
| TI-62 HTB | 15 x 15 x 30 |  | 3 Phase, 415 Volts, 4 Wire 6 KW, | PRICE ON REQUEST | 8,500 |
| TI-62 HTC | 20 x 20 x 30 |  | 3 Phase, 415 Volts, 4 Wire, 7.5 KW, | PRICE ON REQUEST | 8,500 |
| TI-62 HTD | 22.5 x 22.5 x 45 |  | 3 Phase, 415Volts, 4 Wire, 10.7 KW, | PRICE ON REQUEST | 10,000 |

OPTIONALS :-

| 1 | Data logger used for interfacing RS 485 compatible PID controller for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface with 16 x 2 LCD,4 programmable key & status LED’s - PID controller interface - Password protected setting - Data storage capacity – 5000 Readings for each channel for Printer & PC - Programmable logging interval (Print frequency) 1 – 255 min. - Printing Option – Online printing, Date wise printing, Page wise printing - Direct printer (EPSON LX300+) Interface - Memory full (> 70%) Alarm | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485 multi dropping monitor 8 chambers to one software. - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |

SPARE PARTS KEEP IN STOCK

|  |  | LIST PRICE (Amt. in Rs.) |
| --- | --- | --- |
| 1 | Heaters for TI-62 HTA (8 nos. Heater) | PRICE ON REQUEST |
| 2 | Heaters for TI-62 HTB (SPL) (8 nos. Heater) | PRICE ON REQUEST |
| 3 | Heaters for TI-62 HTB (12 nos. Heater) | PRICE ON REQUEST |
| 4 | Heaters for TI-62 HTC (12 nos. Heater) | PRICE ON REQUEST |
| 5 | Heaters for TI-62 HTD (18 nos. Heater) | PRICE ON REQUEST |
| 6 | Thyristor for TI-62 HTA (Single Phase) | PRICE ON REQUEST |
| 7 | Thyristor for TI-62 HTB (SPL) (Single Phase) | PRICE ON REQUEST |
| 8 | Thyristor for TI-62 HTB (Three Phase) | PRICE ON REQUEST |
| 9 | Thyristor for TI-62 HTC (Three Phase) | PRICE ON REQUEST |
| 10 | Thyristor for TI-62 HTD (Three Phase) | PRICE ON REQUEST |
| 11 | Digital Ammeter for TI-62 HTA (Single Phase) | PRICE ON REQUEST |
| 12 | Digital Ammeter for TI-62 HTB (SPL) (Single Phase) | PRICE ON REQUEST |
| 13 | Digital Ammeter for TI-62 HTB (Three Phase) | PRICE ON REQUEST |
| 14 | Digital Ammeter for TI-62 HTC (Three Phase) | PRICE ON REQUEST |
| 15 | Digital Ammeter for TI-62 HTD (Three Phase) | PRICE ON REQUEST |
| 16 | R Type Thermocouple | PRICE ON REQUEST |
| 17 | PID Controller (Main Controller) | PRICE ON REQUEST |
| 18 | Safety Controller (Power and safety contactors) | PRICE ON REQUEST |
| 19 | Safety digital temperature controller with Thermocouple | PRICE ON REQUEST |

- INSULATION CRACKING

The insulation material in this furnace is susceptible to surface cracking as a result of temperature cycling. This is

a normal occurrence and such cracking is not detrimental to the performance of the overall life of the furnace,

under normal operation.

- LOADING THE FURNACE

When heating large object, in particular poor conductors, avoid shielding the thermocouple from the heating

element. Also ensure that nothing is within 15 mm of the elements.

The thermocouple is intended to sense the temperature near the heating element.

However if a large object is placed in the chamber it may record the average Temperature of the object and the

elements, which can lead to over heating of the Elements.

Allow large objects to gain heat at a lower temperature and then reset the controller to a Temperature close to the

desired maximum.

Take care that nothing hits the elements when loading and unloading.

- OPENING THE DOOR

Take great care when loading or unloading the furnace chamber. See section 4.2 before removing a hot object

from the furnace; ensure a suitable surface is available on which to put it.

Do not open the door at high temperature. If possible, do the open at above 2000C. If it is necessary to load or

unload work at elevated temperatures, keep the door open for short a period as possible. The insulation cools

quickly and may crack through thermal shock.

The heating elements and the insulation are very susceptible to mechanical shock. At all times operate the door

gently to avoid mechanical shock.

Opening the door when the furnace is very hot can cause hot air to be sucked by the case cooling fans towards

the thermal cut out and activate it (see section 4.11). Avoid excessive door opening and shut it as soon as

possible.

WARRANTY:

- Heating Elements and Alumina Tube are not warranted.

- All the components of furnace are warranted for the duration of 12 Months against any manufacturing defect.

- The defected component will be replaced.

- The furnace should be used as per the instructions and training.

- Any damage occurring due to mishandling and wrong operational practice is not covered under warranty.

- In case of Quartz Tube Furnace, devitrification of quartz tube occurring from prolong use is not covered in Warranty.

FURNACE - HIGH TEMPERATURE 17000C :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Maximum Temp. Range | 17000C |
| Continuous working Temp. | 16000C |
| Body Construction | MS fabricated with duly powder coated, dual body structure. Control panel will be located at the bottom of the furnace |
| Door | Leak proof Swing aside type door with door safety switch |
| Insulation | Low energy consumption vacuum formed fiber boards, triple wall insulation, Hot wall insulation will be 18000C HD grade, |
| CONTROL SYSTEM |  |
| Heating Rate | Up to 10000C 50C /Min & above 10000C to target temp 50C /Min |
| Heating Element | Kanthal Super Equivalent U shaped MoSi2 Heating elements |
| Skin Temperature | Less than 600C at 16000C |
| Heating Rate | Up to 10000C 50C /Min & above 10000C to target temp 50C /Min |
| Thermocouple (DUAL) | R type with alumina sheath (Main Controller 1 no. & Safety Controller 1 no.) |
| Temperature Controller | Digital Temperature Programmable PID controller with 18 Programs and 8 segments in each program, 1 No. |
| Safety Controller | Re70 safety controller (With Separate Thermocouple as above) |
| Temperature Accuracy | +/- 20C |
| Temperature Uniformity | +/- 150C |
| Thyristor Unit | Thyristor power controller with soft starting facility & current limit facility. |
| Control Switch | Rocker switch, to switch control on/off |
| Safety Switch | Rocker switch, to switch heater supply on/off |
| Power Transformer | Heavy-duty step-down transformer will be provided |
| Control Transformer | 415V/230V, Control transformer will provide for isolation and control circuit safety |
| MCB | MCB will be provided for over current and short circuit protection to the circuit. Also works as main incomer of the system |
| Ammeters | Digital ammeter to indicate load current |
| Indication | Power and safety contactors |
| Contactors | Indicators to indicate load condition |
| SALIENT FEATURES |  |
| - Rugged and durable construction |  |
| - Outer casing is made up of thick mild steel with duly powder coated, gives long service life with extremely resistant. |  |
| - Multilayer insulation with vacuum formed ceramic fibre or refractory bricks with low thermal mass |  |
| - Parallel guided swing aside door design with safety switch. |  |
| - Furnace design with powerful heating elements for providing fast heating rate. |  |
| - Fast heat up and cool down rates |  |
| - Better Temperature Uniformity |  |
| - Temperature Controller with PID Controllers. |  |
| - Power control with Thyristor units for very precise temperature Control, wear-free and noiseless. |  |
| - High mechanical durability, short heating cycle and low power consumption. |  |

| Standard Model | Internal size H x W x D in cm | Outer size H x W x D in cm | Rating KW, Supply & Heater | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-63 HTA | 12.5 x 12.5 x 25 |  | 1 Phase,230 Volts 2 Wire, 3.6 KW | PRICE ON REQUEST | 7,500 |
| TI-63 HTB | 15 x 15 x 20 |  | 1Phase, 230 Volts 2 Wire, 4.1 KW | PRICE ON REQUEST | 8,500 |
| TI-63 HTC | 20 x 20 x 20 |  | 1 Phase,230 Volts 2 Wire 5.06 KW | PRICE ON REQUEST | 9,000 |
| TI-63 HTD | 15 x 15 x 30 |  | 1 Phase,230 Volts 2 Wire, 5.5 KW | PRICE ON REQUEST | 9,000 |
| TI-63 HTE | 20 x 20 x 30 |  | 3 Phase, 415 Volts, 4 Wire, 7.5 KW | PRICE ON REQUEST | 9,000 |
| TI-63 HTF | 30 x 30 x 30 |  | 3 Phase, 415 Volts, 4 Wire, 10 KW | PRICE ON REQUEST | 9,500 |
| TI-63 HTG | 22.5 x 22.5 x 45 |  | 3 Phase, 415 Volts, 4 Wire, 12 KW | PRICE ON REQUEST | 12,000 |
| TI-63 HTH | 30 x 30 x 45 |  | 3 Phase, 415 Volts, 4 Wire, 15 KW | PRICE ON REQUEST | 14,000 |

SPARE PARTS KEEP IN STOCK

| Sr. No. | SPARE PARTS | List Price (Amt. in Rs.) |
| --- | --- | --- |
| 1 | MoSi2 Heaters for TI-63 HTA (4 nos. Heaters) | PRICE ON REQUEST |
| 2 | MoSi2 Heaters for TI-63 HTB (4 nos. Heaters) | PRICE ON REQUEST |
| 3 | MoSi2 Heaters for TI-63 HTC (4 nos. Heaters) | PRICE ON REQUEST |
| 4 | MoSi2 Heaters for TI-63 HTD (6 nos. Heaters) | PRICE ON REQUEST |
| 5 | MoSi2 Heaters for TI-63 HTE (6 nos. Heaters) | PRICE ON REQUEST |
| 6 | MoSi2 Heaters for TI-63 HTF (6 nos. Heaters) | PRICE ON REQUEST |
| 7 | MoSi2 Heaters for TI-63 HTG (6 nos. Heaters) | PRICE ON REQUEST |
| 8 | MoSi2 Heaters for TI-63 HTH (6 nos. Heaters) | PRICE ON REQUEST |
| 9 | Thyristor Unit (Single Phase) for TI-63 HTA,B, C & D | PRICE ON REQUEST |
| 10 | Thyristor Unit (Three Phase) for TI-63 HTE, F, G & H | PRICE ON REQUEST |
| 11 | Digital Ammeter (Three Phase) for TI-63 HTA,B, C & D | PRICE ON REQUEST |
| 12 | Digital Ammeter (Three Phase) for TI-63 HTE, F, G & H | PRICE ON REQUEST |
| 13 | B Type thermocouple | PRICE ON REQUEST |
| 14 | PID Controller (Main Controller) | PRICE ON REQUEST |
| 15 | Safety Controller | PRICE ON REQUEST |
| 16 | Safety digital temperature controller with Thermocouple | PRICE ON REQUEST |

- INSULATION CRACKING

The insulation material in this furnace is susceptible to surface cracking as a result of temperature cycling. This is

a normal occurrence and such cracking is not detrimental to the performance of the overall life of the furnace,

under normal operation.

- LOADING THE FURNACE

When heating large object, in particular poor conductors, avoid shielding the thermocouple from the heating

element. Also ensure that nothing is within 15 mm of the elements.

The thermocouple is intended to sense the temperature near the heating element.

However if a large object is placed in the chamber it may record the average

Temperature of the object and the elements, which can lead to over heating of the Elements.

Allow large objects to gain heat at a lower temperature and then reset the controller to a

Temperature close to the desired maximum.

Take care that nothing hits the elements when loading and unloading.

- OPENING THE DOOR

Take great care when loading or unloading the furnace chamber. See section 4.2 before removing a hot object

from the furnace; ensure a suitable surface is available on which to put it.

Do not open the door at high temperature. If possible, do the open at above 2000C. If it is necessary to load or

unload work at elevated temperatures, keep the door open for short a period as possible. The insulation cools

quickly and may crack through thermal shock.

The heating elements and the insulation are very susceptible to mechanical shock. At all times operate the door

gently to avoid mechanical shock.

Opening the door when the furnace is very hot can cause hot air to be sucked by the case cooling fans towards

the thermal cut out and activate it (see section 4.11). Avoid excessive door opening and shut it as soon as

possible.

WARRANTY:

- Heating Elements and Alumina Tube are not warranted.

- All the components of furnace are warranted for the duration of 12 Months against any manufacturing defect.

- The defected component will be replaced.

- The furnace should be used as per the instructions and training.

- Any damage occurring due to mishandling and wrong operational practice is not covered under warranty.

- In case of Quartz Tube Furnace, devitrification of quartz tube occurring from prolong use is not covered in Warranty.

VERTICAL  LABORATORY AUTOCLAVE  (Regular Model)  :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | Vertical double walled design, has single chamber for Steam & water - Lid. flange & bottom sheet also made of S.S all joints Argon welded. - Joint less silicon Gasket. Heavy duty industrial flange heater. - Lid fitting Pressure gauge 0-30 PSI, safety spring loaded Pressure valve, steam Release valve. |
| Features | Foot lifting arrangement is provided in All models |
| Basket | S. S. Rod Basket |
| Safety | Safety high pressure release valve |
| Vacuum Breaker | Vacuum breaker & self-purging system |
| Supply | 230V AC, Single Phase, 50Hz (TI-121 A to TI-121 G) 415V AC, Three Phase, 50 Hz (TI-121 H) |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 mirror finish. Outer SS 304 mirror finish. |
| GMP Model | Inside S.S. 316 mirror finish. Outer SS 304 mirror finish. |
| Temp. Range | 1210C to 1250C factory set at 1210 C |
| Pressure Range | 15 to 22 PSI, factory set at 15 PSI |

| Models | Capacity Liters | Internal size Dia. x Ht. cm | Rating (KW) | Standard Model (Amt. in Rs.) | GMP Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| TI-121 A | 22 | 25 x 45 | 1.5 | 41,000 | 43,500 | 3,000 |
| TI-121 B | 35 | 30 x 50 | 2.0 | 42,900 | 46,000 | 3,500 |
| TI-121 C | 53 | 35 x 55 | 3.0 | 53,200 | 58,100 | 4,000 |
| TI-121 D | 95 | 45 x 60 | 4.0 | 73,200 | 80,600 | 4,000 |
| TI-121 E | 119 | 45 x 75 | 4.5 | 81,200 | 90,500 | 6,000 |
| TI-121 F | 178 | 55 x 75 | 6.0 | 1,01,500 | 1,13,600 | 8,000 |
| TI-121 G | 225 | 55 x 95 | 6.0 | 1,28,800 | 1,46,900 | 10,000 |
| TI-121 H (THREE PHASE) | 268 | 60 x 95 | 9.0 | 1,50,600 | 1,64,700 | 12,000 |

NOTE : Aavalable with 30 PSI ; will be charges extra 15% in above list price.

| Validation & Documentation | : | 15 OR 30 Minutes Loaded or unloaded condition on one cycle (4 Temp. Probes) along with IQ, OQ & PQ Documents. |
| --- | --- | --- |
| Cost of Validation | : | EXTRA Rs. 15,000/- + GST @ 18% Extra |

| SEMI AUTOMATIC SYSTEM FOR AUTOCLAVE : - Supplied with Microprocessor Based Digital Temperature controller with inbuilt Timer - Low water level cut off system with magnetic float switch for heater’s safety with alarm. | 16,600 |
| --- | --- |
| FULLY AUTOMATIC SYSTEM FOR AUTOCLAVE : - Supplied with Microprocessor Based Digital Temperature controller with inbuilt Timer - Automatic purging of steam with solenoid valve. - Low water level cut off system with magnetic float switch for heater’s safety with alarm. | 23,500 |

OPTIONALS :-

| 1 | Temperature Controller & Timer with Data logger to use for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface - Password protected setting - Programmable logging interval (Print frequency) 1 – 300 Sec. - Direct printer (EPSON LX300+) Interface - F0 Value Calculation & Indication | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485. (To be ordered with Optional no. 1) - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Scanner 4 point temperature scanner complete with sensor. | 33,000 |
| 4 | Digital Pressure indicator (To be ordered with Optional no. 1) | 12,000 |
| 5 | Safety digital temperature controller | 6,000 |
| 6 | Safety pressure control switch | 8,000 |
| 7 | Glass Water level Indicator | 6,300 |

Requirements for Installation :-

- PC requirement (if Software ordered) : 4 GB RAM, Windows 7 or 8,  CD ROM, 2 Nos. USB Port.

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Use of Soft water or D. M. water for humidity formation is recommended. Use of hard or salty water will

damage the boiler heater.

- Do not test any product in the chamber, which is capable of generating combustible mixtures

SINGLE  LEVER  FULLY AUTOMATIC LABORATORY  AUTOCLAVE

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | - Vertical double walled design, has single chamber for Steam & water. - Top lid with Single lever locking facility. - Insulation wound on inner chamber. - Lid. Flange & bottom sheet also made of S.S all joints Argon welded. - Joint less silicon Gasket. Heavy duty industrial flange heater. - Lid fitting Pressure gauge 0-30 PSI, safety spring loaded Pressure valve, steam Release valve. - Fiber/PU Wheel with front lockable for easy movement. |
| Basket | S. S. Rod Basket |
| Safety | - Safety high pressure release valve. - Low water level cut off system with magnetic float switch for heater’s safety with Alarm. - Additional Safety digital temperature controller, in case temperature goes above safety limit hater will be cut off. - Safety Switch – Heater will not energies if lid is not locked properly. - Pressure Interlock on door - Sensor open alarm Visual - Effortless lifting of lid with assisted support (will rest on 450) |
| Vacuum Breaker | Vacuum breaker & self-purging system |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| STD Model | Inside S.S. 304 mirror finish. Outer G.I. Powder coated. |
| GMP Model | Inside S.S. 304 mirror finish. Outer S. S. 304 mirror finish |
| Temp. Range | 121o C to 125o C factory set at 1210 C |
| Pressure Range | 15 to 22 PSI, factory set at 15 PSI |
| Controlling System | - Supplied with Microprocessor Based Digital Temperature controller with inbuilt Timer - Automatic purging of steam with solenoid valve. - Low water level cut off system with magnetic float switch for heater‟s safety with alarm. |
| Safety controller | Safety digital temperature controller |

| Models | Total Internal Size Dia. x Ht. cm | Working Size & Chamber Capacity | Carrier | STD Model (Amt in Rs.) | GMP Model (Amt in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| TI-122 B | 30 x 60 | 30 x 50 / 35 (L) | 25 x 45 | 1,45,000 | 1,58,800 | 3,200/- |
| TI-122 C | 35 x 65 | 35 x 55 / 53 (L) | 30 x 50 | 1,77,500 | 2,02,500 | 3,800/- |
| TI-122 D | 40 x 70 | 40 x 60 / 75 (L) | 35 x 55 | 1,85,000 | 2,11,300 | 4,300/- |
| TI-122 E | 45 x 81 | 45 x 71 /113 (L) | 40 x 33 – 2 Nos. | 2,36,300 | 2,58,100 | 5,400/- |
| TI-122 F | 55 x 87 | 55 x 75 / 180 (L) | 50 x 35 – 2 Nos. | 2,75,000 | 3,06,300 | 6,200/- |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

| Documentation | : | We will provide IQ, OQ & PQ Documents. |
| --- | --- | --- |
| Validation | : | 15 OR 30 Minutes Loaded or unloaded condition on one cycle (4 Temp. Probes) |
| Cost of Validation | : | EXTRA Rs. 15,000/- + GST @ 18% Extra |

OPTIONALS :-

| 1. | Data logger to use for storing data at required interval for printing and / or send to PC for maintaining data record. 1. Simple user interface 2. Password protected setting 3. Programmable logging interval (Print frequency) 1 – 300 Sec. 4. Direct printer (EPSON LX300+) Interface 5. F0 Value Calculation & Indication | 22000/- |
| --- | --- | --- |
| 2. | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485. ( To be ordered with Optional no. 1) - Supports Multiple equipment on single installation - Auto start-up on PC power-up - Access control according to Authority level - Automatic password expiry - Data log report, Alarm report, Audit trail report. | 20,000/- |
| 3. | Digital Pressure indicator. (To be ordered with optional 1) |  |
| 4. | Scanner 4 point temperature scanner complete with sensor. | 33,000 |
| 5. | Safety pressure control switch | 8,000/- |

Single lever autoclave - https://www.youtube.com/watch?v=oy2FQUPeihA

Requirements for Installation :-

            - PC requirement (if Software ordered) : 4 GB RAM, Windows 7 or 8,  CD ROM, 2 Nos. USB Port.

            - Place the chamber where ventilation is good.

            - The chamber must be operated in an ambient of 220C ~ 300C.

            - Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

            - Place the chamber horizontally leveled.

            - Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage

stabilizer (KVA).

            - Proper grounding of the chamber to get satisfactory operation of chamber.

            - Use of Soft water or D. M. water for humidity formation is recommended. Use of hard or salty water will

damage the boiler heater.

9.    Do not test any product in the chamber, which is capable of generating combustible mixtures

VERTICAL  LABORATORY AUTOCLAVE  (PEDAL FREE)  :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | Vertical double walled design, has single chamber for Steam & water. - Unique Pedal free design - Lower Working Height. - Adjustable legs - Lid. flange & bottom sheet also made of S.S all joints Argon welded. - Joint less silicon Gasket. Heavy duty industrial flange heater. - Lid fitting Pressure gauge 0-30 PSI, safety spring loaded Pressure valve, steam - Release valve. |
| Basket | S. S. Rod Basket |
| Safety | Safety high pressure release valve |
| Vacuum Breaker | Vacuum breaker & self-purging system |
| Supply | 230V AC, Single Phase, 50Hz (TI-123 PFA To TI-123 PFG) 415V AC, Three Phase, 50 Hz (TI-123 PFH) |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 mirror finish & Outer SS 304 mirror finish. |
| GMP Model | Inside S.S. 316 mirror finish & Outer SS 304 mirror finish. |
| Temp. Range | 1210C to 1250C factory set at 1210 C |
| Pressure Range | 15 to 22 PSI, factory set at 15 PSI |

| Models | Capacity Liters | Internal size Dia. x Ht. cm | Rating (KW) | Standard Model (Amt. in Rs.) | GMP Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| TI-123 PFA | 22 | 25 x 45 | 1.5 | 46,100 | 51,100 | 3,000 |
| TI-123 PFB | 35 | 30 x 50 | 2.0 | 48,200 | 54,000 | 3,500 |
| TI-123 PFC | 53 | 35 x 55 | 3.0 | 59,400 | 67,700 | 4,000 |
| TI-123 PFD | 95 | 45 x 60 | 4.0 | 82,400 | 94,200 | 4,000 |
| TI-123 PFE | 119 | 45 x 75 | 4.5 | 90,200 | 1,04,600 | 6,000 |
| TI-123 PFF | 178 | 55 x 75 | 6.0 | 1,12,300 | 1,30,800 | 8,000 |
| TI-123 PFG | 225 | 55 x 95 | 6.0 | 1,39,600 | 1,57,800 | 10,000 |
| TI-123 PFH (THREE PHASE) | 268 | 60 x 95 | 9.0 | 1,63,000 | 1,76,400 | 12,000 |

NOTE : Available with 30 PSI ; will be charges extra 15% in above list price.

| Validation & Documentation | : | 15 OR 30 Minutes Loaded or unloaded condition on one cycle (4 Temp. Probes) along with IQ, OQ & PQ Documents. |
| --- | --- | --- |
| Cost of Validation | : | EXTRA Rs. 15,000/- + GST @ 18% Extra |

| SEMI AUTOMATIC SYSTEM FOR AUTOCLAVE : - Supplied with Microprocessor Based Digital Temperature controller with inbuilt Timer - Low water level cut off system with magnetic float switch for heater’s safety with alarm. | 16,600 |
| --- | --- |
| FULLY AUTOMATIC SYSTEM FOR AUTOCLAVE : - Supplied with Microprocessor Based Digital Temperature controller with inbuilt Timer - Automatic purging of steam with solenoid valve. - Low water level cut off system with magnetic float switch for heater’s safety with alarm. | 23,500 |

OPTIONALS :-

| 1 | Temperature Controller & Timer with Data logger to use for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface - Password protected setting - Programmable logging interval (Print frequency) 1 – 300 Sec. - Direct printer (EPSON LX300+) Interface - F0 Value Calculation & Indication | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485. (To be ordered with Optional no. 1) - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Scanner 4 point temperature scanner complete with sensor. | 33,000 |
| 4 | Digital Pressure indicator (To be ordered with Optional no. 1) | 12,000 |
| 5 | Safety digital temperature controller | 6,000 |
| 6 | Safety pressure control switch | 8,000 |
| 7 | Glass Water level Indicator | 6,300 |

Requirements for Installation :-

      - PC requirement (if Software ordered) : 4 GB RAM, Windows 7 or 8,  CD ROM, 2 Nos. USB Port.

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage  stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Use of Soft water or D. M. water for humidity formation is recommended. Use of hard or salty water will

damage the boiler heater.

- Do not test any product in the chamber, which is capable of generating combustible mixtures

TRIPLE WALL SERGICAL AUTOCLAVE  (Foot Lifting OR Pedal Free):-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | Vertical double walled design, has separate chamber for Steam & water. Foot lifting or Unique Pedal free design - Lid. flange & bottom sheet also made of S.S all joints Argon welded. - Joint less silicon Gasket. Heavy duty industrial flange heater. - Lid fitting Pressure gauge 0-30 PSI, safety spring loaded Pressure valve, steam - Release valve. Triple wall construction for Dry Sterilization - Separate tank for water storage (Water Jacket). - Glass Water level indicator with Funnel for water filling. - Separate Drainage for jacket and inner wall - Separate Pressure Guage arrangement - Separate 2 Nos safety valve for External Chamber |
| Basket | S. S. Rod Basket |
| Safety | Safety high pressure release valve |
| Vacuum Breaker | Vacuum breaker & self-purging system |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 mirror finish & Outer S.S. 304 mirror finish. |
| Temp. Range | 1210C to 1250C factory set at 1210 C |
| Pressure Range | 15 to 22 PSI, factory set at 15 PSI |

| Models | Capacity Liters | Internal size Dia. x Ht. cm | Rating (KW) | Foot Lifting Arrangement (Amt. in Rs.) | Pedal Free Arrangement. (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| TI-124 A | 22 | 25 x 45 | 1.5 | 60,500 | 62,900 | 3,000 |
| TI-124 B | 35 | 30 x 50 | 2.0 | 62,700 | 67,700 | 3,500 |
| TI-124 C | 53 | 35 x 55 | 3.0 | 74,900 | 79,300 | 4,000 |
| TI-124 D | 95 | 45 x 60 | 4.0 | 99,800 | 1,02,700 | 4,000 |
| TI-124 E | 119 | 45 x 75 | 4.5 | 1,08,500 | 1,10,800 | 6,000 |
| TI-124 F | 178 | 55 x 75 | 6.0 | 1,37,300 | 1,41,100 | 8,000 |

NOTE : Available with 30 PSI ; will be charges extra 15% in above list price.

| Validation & Documentation | : | 15 OR 30 Minutes Loaded or unloaded condition on one cycle (4 Temp. Probes) along with IQ, OQ & PQ Documents. |
| --- | --- | --- |
| Cost of Validation | : | EXTRA Rs. 15,000/- + GST @ 18% Extra |

| SEMI AUTOMATIC SYSTEM FOR AUTOCLAVE : - Supplied with Microprocessor Based Digital Temperature controller with inbuilt Timer - Low water level cut off system with magnetic float switch for heater’s safety. | 16,600 |
| --- | --- |
| FULLY AUTOMATIC SYSTEM FOR AUTOCLAVE : - Supplied with Microprocessor Based Digital Temperature controller with inbuilt Timer - Automatic purging of steam with solenoid valve. - Low water level cut off system with magnetic float switch for heater’s safety. | 23,500 |

OPTIONALS :-

| 1 | Temperature Controller & Timer with Data logger to use for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface - Password protected setting - Programmable logging interval (Print frequency) 1 – 300 Sec. - Direct printer (EPSON LX300+) Interface - F0 Value Calculation & Indication | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485. (To be ordered with Optional no. 1) - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Scanner 4 point temperature scanner complete with sensor. | 33,000 |
| 4 | Digital Pressure indicator (To be ordered with Optional no. 1) | 12,000 |
| 5 | Safety digital temperature controller | 6,000 |
| 6 | Safety pressure control switch | 8,000 |
| 7 | Glass water level Indicator | 6,300 |

Requirements for Installation :-

- PC requirement (if Software ordered) : 4 GB RAM, Windows 7 or 8,  CD ROM, 2 Nos. USB Port.

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Use of Soft water or D. M. water for humidity formation is recommended. Use of hard or salty water will

damage the boiler heater.

- Do not test any product in the chamber, which is capable of generating combustible mixtures

VERTICAL LABORATORY AUTOCLAVE (RADIAL LOCKING):-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | Vertical double walled design has single chamber for Steam & water. - Top Lid with Radial Locking & foot lifting facility. - Lid is tightened by Radial Locking system fitted through a pedal lifting. - Lid. flange & bottom sheet also made of S.S all joints Argon welded. - Joint less silicon Gasket. Heavy duty industrial flange heater. - Lid fitting Pressure gauge 0-30 PSI, safety spring loaded Pressure valve, steam Release valve. - Hydro test performed at 35 PSI |
| Basket | S. S. Rod Basket |
| Safety | Safety high pressure release valve |
| Vacuum Breaker | Vacuum breaker & self-purging system |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 mirror finish. Outer SS 304 mirror finish. |
| Temp. Range | 1210 C to 1250 C factory set at 1210 C |
| Pressure Range | 15 to 22 PSI, factory set at 15 PSI |

| Models | Capacity Liters | Internal size Dia. x Ht. cm | Rating (KW) | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-125 RLB | 35 | 30 x 50 | 2.0 | 99,015 | 3,500 |
| TI-125 RLC | 53 | 35 x 55 | 3.0 | 1,22,010 | 4,000 |
| TI-125 RLD | 95 | 45 x 60 | 4.0 | 1,48,575 | 4,000 |
| TI-125 RLE | 119 | 45 x 75 | 4.5 | 1,71,465 | 6,000 |
| TI-125 RLF | 178 | 55 x 75 | 6.0 | 2,13,780 | 8,000 |

| SEMI AUTOMATIC SYSTEM FOR AUTOCLAVE : - Supplied with Microprocessor Based Digital Temperature controller with inbuilt Timer - Low water level cut off system with magnetic float switch for heater’s safety. | 16,600 |
| --- | --- |
| FULLY AUTOMATIC SYSTEM FOR AUTOCLAVE : - Supplied with Microprocessor Based Digital Temperature controller with inbuilt Timer - Automatic purging of steam with solenoid valve. - Low water level cut off system with magnetic float switch for heater’s safety. | 23,500 |

NOTE : Available with 30 PSI ; will be charges extra 15% in above list price.

| Validation & Documentation | : | 15 OR 30 Minutes Loaded or unloaded condition on one cycle (4 Temp. Probes) along with IQ, OQ & PQ Documents. |
| --- | --- | --- |
| Cost of Validation | : | EXTRA Rs. 15,000/- + GST @ 18% Extra |

OPTIONALS :-

| 1 | Temperature Controller & Timer with Data logger to use for storing data at required interval for printing and / or send to PC for maintaining data record. - Simple user interface - Password protected setting - Programmable logging interval (Print frequency) 1 – 300 Sec. - Direct printer (EPSON LX300+) Interface - F0 Value Calculation & Indication | 22,000 |
| --- | --- | --- |
| 2 | CFR 21 Part 11 compliance window based communication software for data management. Complete with RS 485. (To be ordered with Optional no. 1) - Password protected - Logger & Scanner report with graph - Online change controller setting - Logger & Scanner Alarm report - Audit Trail report - Backup & Restore Utility - Create user & give the password | 20,000 |
| 3 | Scanner 4 point temperature scanner complete with sensor. | 33,000 |
| 4 | Digital Pressure indicator To be ordered with Optional no. 1 | 12,000 |
| 5 | Safety digital temperature controller | 6,000 |
| 6 | Safety pressure control switch | 8,000 |

Requirements for Installation :-

      - PC requirement (if Software ordered) : 4 GB RAM, Windows 7 or 8,  CD ROM, 2 Nos. USB Port.

      - Place the chamber where ventilation is good.

      - The chamber must be operated in an ambient of 220C ~ 300C.

      - Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

      - Place the chamber horizontally leveled.

      - Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage stabilizer (KVA).

      - Proper grounding of the chamber to get satisfactory operation of chamber.

      - Use of Soft water or D. M. water for humidity formation is recommended. Use of hard or salty water will

damage the boiler heater.

- Do not test any product in the chamber, which is capable of generating combustible mixtures

VERTICAL LABORATORY AUTOCLAVE (PORTABLE) :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | Vertical Single walled design, has single chamber for Steam & water - Lid. flange & bottom sheet also made of S.S all joints Argon welded. - Joint less silicon Gasket. Heavy duty industrial flange heater. - Lid fitting Pressure gauge 0-30 PSI, safety spring loaded Pressure valve, steam Release valve. |
| Features | Lid lifting arrangement |
| Basket | S. S. Rod Basket |
| Safety | Safety high pressure release valve |
| Vacuum Breaker | Vacuum breaker & self-purging system |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 mirror finish. Outer SS 304 mirror finish. |
| Temp. Range | 1210C to 1250C factory set at 1210 C |
| Pressure Range | 15 to 22 PSI, factory set at 15 PSI |

| Models | Capacity Liters | Internal size Dia. x Ht. cm | Rating (KW) | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-121 PA | 18 | 30 x 25 | 1.5 | 19,950 | 1,500 |
| TI-121 PB | 21 | 30 x 30 | 1.5 | 23,100 | 1,800 |

NOTE : Available with 30 PSI ; will be charges extra 15% in above list price.

| Validation & Documentation | : | 15 OR 30 Minutes Loaded or unloaded condition on one cycle (4 Temp. Probes) along with IQ, OQ & PQ Documents. |
| --- | --- | --- |
| Cost of Validation | : | EXTRA Rs. 15,000/- + GST @ 18% Extra |

OPTIONALS :-

| 1 | Imported auto tuned PID Controller cum Indicator with PT-100 Sensor and in build Timer | 7,500 |
| --- | --- | --- |

Requirements for Installation :-

- PC requirement (if Software ordered) : 4 GB RAM, Windows 7 or 8,  CD ROM, 2 Nos. USB Port.

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage   stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Use of Soft water or D. M. water for humidity formation is recommended. Use of hard or salty water will

damage the boiler heater.

- Do not test any product in the chamber, which is capable of generating combustible mixtures

VERTICAL LAMINAR AIR FLOW UNIT

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Filter Details Mini pleat HEPA Filter | MOC- Anodized Aluminium Efficiency- H-14 as per EN1822 (99.997% down to 0.3 micron) Type- Miniplate Gasket- Silicon / Neoprene Media- Micro Glass Fibers |
| Pre Filter | MOC- Anodized Aluminium Efficiency-EU – 5 rating, Efficiency: 95% down to 5 micron Type- Flange type washable Gasket- Neoprene Media- Synthetic Media with HDPE Net, Non-Woven Type |
| Control System | Microprocessor based control panel is located on the equipment with Keypad & LCD - Inbuilt counter for U.V. light - Inbuilt counter for Hepa Filter - Inbuilt sensor with pressure display - Magnehelic guage. |
| Face Velocity | Face Velocity of HEPA is adjustable from 90 ± 10 FPM or 0.45 ± 0.05m/s |
| Air Quality | Satisfy ISO 5,Class 100 requirements as per ISO-14644 &IEST-RP-CC001.3dv |
| Blower motor | Single Phase Type- Backward Curved Centrifugal fans MOC- Aluminium Body. Noise -Dynamically &Statically balanced blower for less vibration &noise level. |
| Instrumentation & Accessories | - U.V Light with Timer - LED Lighting to provide an intensity of 600 Lux - Front sliding door acrylic |
| Door Opening Alarm | Alarm triggers in case front door is raised more than 8” (i.e. safe clear door opening during operation) |
| Magnehelic Gauge | Analog type Gauge of range: 0 to 25 mmwc |
| Electrical and Other Accessories | - On /Off Switch for Motor Blower - Electrical Interlocking between U.V. Light and Florescent light (220V, 5A Single Phase, 50Hz, Lighting additional power point - Levelling Legs - PU Wheels. |
| Standard Model | S.S. 304 Table & Outer G.I. Powder coated |
| GMP Model | S.S. 304 Table & Outer S.S. 304 Matt Finish |

| Models | Working Size W x D x H (in feet) | Outer Size Wx D x H (in feet) | Hepa Filter | Pre Filter | Standard Model (Amt. in Rs.) | GMP Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TI-51 A VLAF | 2 x 2 x 2 | 85 x 75 x 210 | 1 | 1 | 1,06,100 | 1,38,600 | 8,000 |
| TI-51 B VLAF | 3 x 2 x 2 | 85 x 108 x 210 | 1 | 1 | 1,28,700 | 1,77,400 | 10,000 |
| TI-51 C VLAF | 4 x 2 x 2 | 85 x 140 x 210 | 1 | 1 | 1,44,700 | 2,16,300 | 12,000 |
| TI-51 D VLAF | 6 x 2 x 2 |  | 1 | 1 | 2,23,700 | 3,22,900 | 15,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

Below tests or any one test can be performed at your site at extra charges Rs. 30,000/-

            - Velocity Measurement

            - Hepa Integrity test

            - Paticle Count

Documentation :  We will provide IQ, OQ & PQ Documents at the time of Installation.

MOC Certificate of SS 304 and Filter Test Certificate.

Pre-Filter, Hepa Filter, Motor Blower, Material Of Construction & Warranty Certificate.

Warranty     : Filter warranty is not supported after installation of Laminar Air flow.

OPTIONALS :-

| 1 | Digital pressure display with Magnehelic guage & filter failure Alarm | 6,000 |
| --- | --- | --- |
| 2 | Gas Cork arrangement (1/2 inch) | 3,250 |
| 3 | Water Cork arrangement (1/2 inch) | 3,250 |
| 4 | Air Velocity Display – Real time monitoring & display of air velocity with inbuilt digital pressure display. | 33,500 |

Requirements for Installation :-

      - Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 15 Amp). It is recommended to use a voltage stabilizer (KVA).

      - Proper grounding of the chamber to get satisfactory operation of chamber.

      - Surrounding area should be clean & dust free.

      - Please confim the Room & Door size with extranal size of Laminar

HORIZONTAL LAMINAR AIR FLOW UNIT

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Filter Details Mini pleat HEPA Filter | MOC- Anodized Aluminium Efficiency- H-14 as per EN1822 (99.997% down to 0.3 micron) Type- Miniplate Gasket- Silicon / Neoprene Media- Micro Glass Fibers |
| Pre Filter | MOC- Anodized Aluminium Efficiency-EU – 5 rating, Efficiency: 95% down to 5 micron Type- Flange type washable Gasket- Neoprene Media- Synthetic Media with HDPE Net, Non-Woven Type |
| Control System | Microprocessor based control panel is located on the equipment with Keypad & LCD - Inbuilt counter for U.V. light - Inbuilt counter for Hepa Filter - Inbuilt sensor with pressure display – magnehelic guage. |
| Face Velocity | Face Velocity of HEPA is adjustable from 90 ± 10 FPM or 0.45 ± 0.05m/s |
| Air Quality | Satisfy ISO 5,Class 100 requirements as per ISO-14644 &IEST-RP-CC001.3dv |
| Blower motor | Single Phase Type- Backward Curved Centrifugal fans MOC- Aluminium Body. Noise -Dynamically &Statically balanced blower for less vibration &noise level. |
| Instrumentation & Accessories | - U.V Light with timer - LED Lighting to provide an intensity of 600 Lux |
| Magnehelic Gauge | Analog type Gauge of range: 0 to 25 mmwc |
| Electrical and Other Accessories | - On /Off Switch for Motor Blower - Electrical Interlocking between U.V. Light and Florescent light (220V, 5A Single Phase,50Hz, Lighting additional power point. - Levelling Legs. |
| Standard Model | S.S.304 Table & Outer G.I. Powder coated |
| GMP Model | S.S. 304 Table & Outer S.S. 304 Matt Finish |

| Models | Working Size Wx D x H (in feet) | Hepa Filter | Pre Filter | Standard Mode (Amt. in Rs.) | GMP Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| TI-52 A HLAF | 2 x 2 x 2 | 1 | 1 | 1,01,800 | 1,31,100 | 8,500 |
| TI-52 B HLAF | 3 x 2 x 2 | 1 | 1 | 1,23,900 | 1,63,800 | 10,500 |
| TI-52 C HLAF | 4 x 2 x 2 | 1 | 1 | 1,44,100 | 2,07,200 | 12,000 |
| TI-52 D HLAF | 6 x 2 x 2 | 1 | 1 | 1,92,500 | 2,60,700 | 15,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

Below tests or any one test can be performed at your site at extra charges Rs. 30,000/-

- Velocity Measurement

- Hepa Integrity test

- Paticle Count

Documentation :  We will provide IQ, OQ & PQ Documents at the time of Installation.

MOC Certificate of SS 304 and Filter Test Certificate.

Pre-Filter, Hepa Filter, Motor Blower, Material Of Construction & Warranty Certificate.

Warranty        :  Filter warranty is not supported after installation of Laminar Air flow.

OPTIONALS :-

| 1 | Digital pressure display with Magnehelic guage with filter failure alarm | 6,000 |
| --- | --- | --- |
| 2 | Gas Cork arrangement (1/2 inch) | 3,250 |
| 3 | Water Cork arrangement (1/2 inch) | 3,250 |
| 4 | Air Velocity Display – Real time monitoring & display of air velocity with inbuilt digital pressure display | 33,500 |

Requirements for Installation :-

      - Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 15 Amp). It is recommended to use a voltage stabilizer (KVA).

      - Proper grounding of the chamber to get satisfactory operation of chamber.

      - Surrounding area should be clean & dust free.

      - Please confim the Room & Door size with extranal size of Laminar.

BIO SAFETY CABINET  TYEPE  II  A2

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Filter Details Mini pleat HEPA Filter | MOC- Anodized Aluminium Efficiency- H-14 as per EN1822 (99.997% down to 0.3 micron) Type- Miniplate Gasket- Silicon / Neoprene Media- Micro Glass Fibers |
| Pre Filter | MOC- Anodized Aluminium Efficiency-EU – 5 rating, Efficiency: 95% down to 5 micron Type- Flange type washable Gasket- Neoprene Media- Synthetic Media with HDPE Net, Non-Woven Type |
| Control System | Microprocessor based control panel is located on the equipment with Keypad & LCD - Inbuilt counter for U.V. light - Inbuilt counter for Hepa Filter - Inbuilt sensor with pressure display – magnehelic gauge |
| Air Flow Pattern | Unidirectional Vertical down flow 70% re-circulated & 30% exhausted with makeup air coming in as inflow throw table front perforation. Average air flow velocity with uniformly of +20% - Vertical down flow 60 to 90 fpm - Inflow 90 fpm |
| Guidelines | EN 12469 -2000 (European Standard) |
| Blower motor | Single Phase Type- Backward Curved Centrifugal fans MOC- Aluminium Body. Noise -Dynamically & Statically balanced blower for less vibration & noise level. |
| Instrumentation & Accessories : | - U.V Light - LED Lighting to provide an intensity of 600 Lux - Front sliding door acrylic |
| Door Opening Alarm | Alarm triggers in case front door is raised more than 8” (i.e. safe clear door opening during operation) |
| Magnehelic Gauge | Analog type Gauge of range: 0 to 25 mmwc |
| Electrical and Other Accessories | 1. On /Off Switch for Motor Blower 2. Electrical Interlocking between U.V. Light and Florescent light (220V, 5A Single Phase,50Hz, Lighting additional power point. - Levelling Legs. - Castor wheels. |
| STD Model | S.S. 304 Table & Outer GI Powder Coated |
| GMP Model | S.S. 304 Table & Outer S.S. 304 Matt Finish |

| Models | Working Size Wx D x H (in feet) | STD Model (Amt. in Rs.) | GMP Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- |
| TI- 53A - BSC | 2 x 2 x 2 | 2,54,400 | 2,59,600 | 6,000 |
| TI- 53B - BSC | 3 x 2 x 2 | 2,78,500 | 2,84,300 | 7,000 |
| TI- 53C - BSC | 4 x 2 x 2 | 3,02,600 | 3,19,500 | 8,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

Below tests or any one test can be performed at your site at extra charges Rs. 30,000/-

1.  Velocity Measurement

2.  Hepa Integrity test

3.  Paticle Count

Documentation :  We will provide IQ, OQ & PQ Documents at the time of Installation.

MOC Certificate of SS 304 and Filter Test Certificate.

Pre-Filter, Hepa Filter, Motor Blower, Material Of Construction & Warranty Certificate.

Warranty        : Filtter warranty is not supported ofter installation of Laminar Air flow.

OPTIONALS :-

| 1 | Digital pressure display | 6,000 |
| --- | --- | --- |
| 2 | Gas Cork arrangement (1/2 inch) | 3,250 |
| 3 | Water Cork arrangement (1/2 inch) | 3,250 |

Requirements for Installation :-

      - Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 15 Amp). It is recommended to use a voltage

stabilizer (KVA).

      - Proper grounding of the chamber to get satisfactory operation of chamber.

      - Surrounding area should be clean & dust free.

FUME HOOD :

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8421 |
| Construction | Double wall construction with Exterior body of the fume hood is made of mild steel with epoxy powder coated, & Interior body of hood is made POLYPROPYLENE SMOOTH FINISH SHEET OR FRP COATING SHEET. Base with Heavy M.S. Powder coated. |
| Blower Motor | - Blower with Exhaust type system. - Blower with Anti Coregent Coating. - Low Noise level - Fume chamber will be provided with a powerful blower & suitably designed to create necessary suction pressure. |
| Sash (Shutter) | Vertical rising counter- balanced sash. Toughened float glass sash for full work area visibility (5 mm thick), Smooth & light sash operation. Clear opening height (650 mm). |
| Airfoil | Aerodynamic design Horizontal airfoil mounted on the worktop made of thick CRCA sheet. |
| Air Velocity | 0.3 – 0.8 m/s. |
| Exhaust Duct | Fabric flexible pipe of 3 mtr for Exhaust |
| Worktop | Chemical resistant slash & spillage proof dished ‘Jet black granite’ worktop (18 to 20 mm thick) |
| Supply | 230V AC, Single Phase, 50Hz |
| Lighting: | LED light fitting for proper illumination |
| Electrical utilities | 2 nos. of 5/15 amp electrical points shall be provided on the right side of the fume hood with tube light switch. Pushbutton starter & electrical connector shall be fixed on the top of the fume hood |
| Utility Services | Remote operated services fitting will be provided inside the fume hood with 2 nos. for Gas |

| Model | Internal Size H x W x D in cm | External size H x W x D in cm | Motor Position | With PP Coating (Amt. in Rs) | With FRP Coating (Amt. in Rs) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| TABLE TOP MODEL |  |  |  |  |  |  |
| TI-12 FMA | 110 x 60 x 75 | 160 x 92 x 112 | Inbuilt Blower Motor | 88,300 | 1,02,400 | 12,000 |
| TI-12 FMB | 120 x 90 x 75 | 225 x 122 x 112 | Inbuilt Blower Motor | 1,00,800 | 1,18,900 | 14,000 |
| FLOOR MODEL - STORAGE CABINET WITH WHEEL |  |  |  |  |  |  |
| TI-12 FMC | 120 x 120 x 75 | 295 x 155 x 112 | Inbuilt Blower Motor | 1,60,100 | 1,87,400 | 22,000 |
| TI-12 FMD | 120 x 150 x 75 | 295 x 185 x 112 | Inbuilt Blower Motor | 1,78,000 | 2,14,600 | 26,000 |

***  Please Confirm Room Door Size (for equipment entry) and ceiling height.

Erection, Installation :

Erection, installation & ducting charges for the fume hood will be charged extra.

Validation :

- Fume Hood Validation (Optional – Charged Extra):

  - Air Velocity Test:
Conducted at 25 cm sash opening, with face velocity maintained at 100 FPM (± 30 FPM).

  - Measurement Points:
Air velocity readings will be taken at 6 different positions across the working area.

OPTIONALS :-

| 1 | Gas line with Tap | 6,200 |
| --- | --- | --- |
| 2 | Sink for water with drain tap facility | 10,000 |
| 3 | HDPE pipe for ducting | 2,850 per mtr. |
| 4 | Temprature controller to view temperature inside fume hood or test sample with flexible sensor.(measure tempratire from -50°C to 300°C) | 6,000 |
| 5 | Storage Cabinet for TI-12 FMA & TI-12 FMB - Used for storage of small lab instruments and mild acids. - Chemical storage base cabinet acting as a support stand for fume hood with doors & locking arrangements. - Fiber/PU Wheel with front lockable for easy movement | 25,000 (Packing Charges 2,500 extra) |

Note :

Please check the fume hood outer dimensions against room size and door entrance for proper installation.

Requirements for Installation :

- Leave minimum 1.5 ft.space from the sides and leave 1.5 ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 32 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Provide wall openings / core cutting for ducting in costomer scope

COD DIGESTOR  :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Construction | - Double walled construction. Complete Equipment made from CRCA Steel, duly powder coat finished. - Highly insulated with compressed Cerawool between the walls to minimize the heat dissipation. - Aluminium Block having holes used to hold Glass tubes & is heated to maintain temperature up to 180 °C. |
| Timer Range | 999 Minutes. |
| Block Hole Dimension | 40 mm Ø X 80 mm Depth. |
| Glass Tube Size | 38 mm Ø. |
| Air Condenser | 600 mm Height. |
| Accessories | Available With Below Accessories. - Glass Tubes - Air Condensers - Tube Stand - Condenser Stand |
| Supply | 230 Volts AC Single phase 50 Hz. |

| SPECIFICATIONS |  |
| --- | --- |
| Temperature Range | 50C Above ambient to 1800C. |
| Temperature Accuracy | + 10C. |
| Control System | Dual Display Microprocessor based PID Temprature Controller cum Timer with & PT -100 sensors. |

| Models | Capacity | Heater Ratings | No of tubes & Air condenser | (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-101 | 06 Nos. (3 x 2 Rows) | 400 Watts. | 06 | 27,100 | 2,000 |
| TI-102 | 09 Nos. (3 x 3 Rows) | 500 Watts. | 09 | 36,200 | 2,500 |
| TI-103 | 15 Nos. (5 x 3 Rows) | 750 Watts. | 15 | 48,200 | 3,500 |
| TI-104 | 24 Nos. (6 x 4 Rows) | 1200 Watts. | 24 | 67,100 | 5,000 |

OPTIONALS :-

| 1 | S.S. Cooling Bath | 3,000 |
| --- | --- | --- |

SPARES

| 2 | Air Condensers | 1,100 |
| --- | --- | --- |
| 3 | Reaction Vessel | 800 |

STATIC PASS BOX

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8421 |
| Certification | CE Certification |
| Type | Wall and floor Mounted |
| Construction | SS 304 Inner and Outer Matt finish with SS 304 frame for support to Pass Box. |
| Controller | Micro controller based LCD |
| Interlocking of doors | - Two doors are interlocked with Magnet. - Toughend Glass on both side of door to for view inside portion. |
| Features | - LED Light intensity more than 500 lux. - V Light.(Interlocking with LED light) for Germicidal Activity - Buzzer Indication to know Material KEPT inside - Indication Lamp to know door locked or unlocked |
| Supply | 230V, 50hz Frequency, |

| Models | Internal Size H x W x D in cm | External Size H x W x D in cm | PRICE (Amt. In Rs.) | Packing Charges |
| --- | --- | --- | --- | --- |
| TI-80 A SPB | 45 x 45 x 45 | 65 x 57 x 52 | 70,200 | 1,500 |
| TI-80 B SPB | 60 x 60 x 60 | 81 x 73 x 68 | 77,200 | 3,000 |
| TI-80 C SPB | 75 x 75 x 75 | 95 x 87 x 82 | 94,500 | 3,500 |
| TI-80 D SPB | 90 x 90 x 90 | 110 x 103 x 99 | 1,31,300 | 4,500 |

Documentation : 1) We will provide IQ, OQ & PQ Documents at the time of Installation.

2)  Standard operating manual (User Manual) & Material of Construction.

OPTIONALS :-

| 1 | S.S. 304 Stand to Support wall mounting pass box (1 nos.) | Price |
| --- | --- | --- |
|  | TI-80 A SPB | 3,400 |
|  | TI-80 B SPB | 4,200 |
|  | TI-80 C SPB | 5,000 |
|  | TI-80 D SPB | 6,700 |

Requirements for Installation :-

- Required cut-out in puf panel OR wall as per the external size(<1cms than external size) of pass box.

- Required electric power supply 230 volt 50Hz.

DYNAMIC PASS BOX

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8421 |
| Certification | CE Certification |
| Construction | SS 304 Inner and Outer with SS 304 frame for support to Pass Box. |
| Controller | LCD Display with ultra-touch keypad |
| Interlocking of doors | - Two doors are interlocked with Magnet to avoid cross contamination. - Glass on both side of door to view inside portion. |
| Filtration | 1st Stage- Fine Pre Filter - Flange type - Efficiency:-95 %down to 5 micron - Grade- EU-5 2nd stage- HEPA Filter of 99.997%down to 0.3 micron - Box Type - Efficiency:- 99.997% Down to 0.3 Micron - Grade- EU13 |
| Motor Blower | Dynamically & Statically balanced blower to less vibration & noise level. 220-230Volts , 50Hz, Single Phase. |
| Velocity | 0.45m/s +/- 0.05mps |
| Air Quality | Class 100 |
| Noise | Less than 67 db |
| Features | - LED Light intensity more than 500 lux. - UV Light.(Interlocking with LED light) - Inbuild Hr Mtr - Buzzer Indication to know Material KEPT inside - Indication Lamp to know door locked or unlocked - Toughened glass door |

| Models | Internal Size H x W x D in cm | External Size H x W x D in cm | PRICE Amt. in Rs. | Packing Charges |
| --- | --- | --- | --- | --- |
| TI-81 A DPB | 45 x 45 x 45 | 90 x 57 x 52 | 1,19,100 | 2,000 |
| TI-81 B DPB | 60 x 60 x 60 | 110 x 73 x 68 | 1,30,900 | 4,000 |
| TI-81 C DPB | 75 x 75 x 75 | 125 x 87 x 82 | 1,58,100 | 6,000 |
| TI-81 D DPB | 90 x 90 x 90 | 141 x 103 x 99 | 1,94,500 | 7,000 |

Documentation   :  We will provide IQ, OQ & PQ Documents at the time of Installation.

Standard operating  manual (User Manual), Hepa Filter, Motor Blower, MOC,

OPTIONALS :-

| 1 | S.S. 304 Stand to Support wall mounting pass box (1 nos.) | Price |
| --- | --- | --- |
|  | TI-81 A DPB | 4,200 |
|  | TI- 81 B DPB | 5,000 |
|  | TI-81 C DPB | 6,000 |
|  | TI-81 D DPB | 7,000 |

Requirements for Installation :-

- Required cut-out in puf panel OR wall as per the external size(<1cms than external size) of pass box.

- Required electric power supply 230 volt 50Hz.

LABORATORY HOT PLATE (Circular)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Standard Model | A thick gauge Mild Steel (M.S.) body & Cast Iron Top |
| Top | A thick well machined top. This top is grooved from bottom to circulate spiral heaters. |
| Heaters | Heating element is housed into the grooves with porcelain beads underneath of the top. |
| Features | Tempo only manufacturer using grooved top with best quality for uniform heating. Hot plates manufactured according to the industries standards to suit very high corrosive atmosphere and ensured long life performance. |
| Dimension | 20.5 cm. |
| Rating | 1.0 kw. |

| Models | Max. Temperature | Temperature Control | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- |
| TI-85 ER | 3500C | Energy Regulator | 3,800 | 600 |
| TI-85 TC | 3000C | Thermostatic Controller | 4,100 | 600 |

| Optional | Imported auto tuned PID Controller cum Indicator with K type Sensor | 5,600 |
| --- | --- | --- |

LABORATORY HOT PLATE (Rectangular)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Standard Model | A thick gauge Mild Steel (M.S.) body & Cast Iron Top |
| S S 304 & MS Body | S.S 304 Top & Body MS Powder Coated |
| S S 304 Top & Body | Top & Body Both S.S 304 |
| Top | A thick well machined top. This top is grooved from bottom to circulate spiral heaters. |
| Heaters | Heating element is housed into the grooves with porcelain beads underneath of the top. |
| Temperature | Maximum surface Temperature 300 0C. |
| Features | Tempo only manufacturer using grooved top with best quality for uniform heating. Hot plates manufactured according to the industries standards to suit very high corrosive atmosphere and ensured long life performance. |

| Models | Dim. In cm. | Rating KW | Temperature Controller | Standard Model (Amt. in Rs.) | S. S 304 Top & Body MS (Amt. in Rs.) | S.S 304 Top & Body (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TI-70 ERA | 25 x 40 | 1.2 | Energy Regulator | 8,900 | - | - | 600 |
| TI-70 ERB | 30 x 45 | 2.0 | Energy Regulator | 11,800 | - | - | 700 |
| TI-70 ERC | 45 x 60 | 3.0 | Energy Regulator | 21,600 | - | - | 900 |

| TI-70 TCA | 25 x 40 | 1.2 | With Thermostat | 9,200 | 11,100 | 14,200 | 600 |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TI-70 TCB | 30 x 45 | 2.0 | With Thermostat | 12,100 | 13,200 | 18,400 | 700 |
| TI-70 TCC | 45 x 60 | 3.0 | With Thermostat | 21,900 | 25,200 | 29,400 | 900 |

| Optional | Imported auto tuned PID Controller cum Indicator with K type Sensor | 5,600 |
| --- | --- | --- |

MAGNETIC STIRRER WITH HOT PLATE :

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Construction | TEMPO Magnetic Stirrers with Aluminum top housing are most compact models. Permanent Magnet DC (PMDC) motor gives high torque even at lower speeds and maintains speed stability despite viscosity or volume changes Diﬀerent models available for all types of laboratory stirring applications up to 5 liters Accurate steeples speed control with digital display y allows smooth variation up to 1200 rpm. Heat energy in controlled by energy regulator |
| Supply | 230V AC, Single Phase |
| Maximum Temperature | 220° C |

| Models | Max. Speed (RPM) | Stirring Capacity (Ltrs.) | Stirring Pedal Dimension (Dia. x Length) | Heating Capacity (watt.) | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| TI–72-1MLH | 1200 | 1 | 9 x 25 | 150 | 9,300 | 800 |
| TI–72-2MLH | 1200 | 2 | 9 x 35 | 300 | 10,000 | 900 |
| TI–72-5MLH | 1200 | 5 | 13 x 50 | 500 | 12,600 | 1000 |

| Optional | Digital Temperature Controller | 6,600 |
| --- | --- | --- |

MAGNETIC STIRRER WITH OUT HOT PLATE :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 84798200 |
| Certification | CE Certification |
| Construction | - TEMPO Magnetic Stirrers with Aluminum top housing are most compact models. - Permanent Magnet DC (PMDC) motor gives high torque even at lower speeds and maintains speed stability despite viscosity or volume changes. - Diﬀerent models available for all types of laboratory stirring applications up to 5 liters - Accurate steeples speed control with digital display y allows smooth variation up to 1200 rpm. - Heat energy in controlled by energy regulator |
| Supply | 230V AC, Single Phase |

| Models | Max. Speed (RPM) | Stirring Capacity (Ltrs.) | Stirring Pedal Dimension (Dia. x Length) | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-72 1ML | 1200 | 1 | 9 x 25 | 7,500 | 800 |
| TI-72 2ML | 1200 | 2 | 9 x 35 | 7,700 | 900 |
| TI-72 5ML | 1200 | 5 | 13 x 50 | 9,900 | 1000 |

SAILENT FEATURES :

- PMDC motor for higher torque even at low speeds.

- Better speed regulation even with small volume & low speed.

- Accurate steeples speed control maintains excellent speed stability.

- Digital speed indicate or for display of stirring speed.

- Totally enclosed unit

CRYO BATH :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | Double walled construction with inner S.S. argon welded one piece pot which avoids leakages. |
| Insulation | 2” Thick PUF Insulation ensures temperature stability & reduced energy consumption. |
| Water Circulation | Circulation pump is provided for better uniformity of Temperature instead of stirrer. Side mounted circulating pump 18 ltr. /min capacity, single phase, 230 Volts. |
| Heating System | S.S. Tubular heater (Water immersion type) is used at bottom for better heat distribution. |
| Lid | Top mounted removable lid. |
| Refrigeration system | CFC Free hermitically sealed Emerson Copeland make compressor for better cooling with R 134 A (Eco friendly) refrigerant with time delay to safe guard compressor. |
| Safety | - Compressor ON delay timer (2 min.) to safe guard the compressor - In built over load protector provided for hermetically sealed compressor. |
| Features | - Fiber/PU Wheel with front lockable for easy movement. - Circulation pump use for Drain (Do not dry run circulation pump). |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 argon welded one piece pot. Outer Galvanized Iron (GI) sheet with powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 argon welded one piece pot. Outer S.S. 304. |
| G.M.P. Model | Inside S.S. 316 argon welded one piece pot. Outer S.S. 304. |
| Temperature Range | 50 C to 900 C |
| Temperature Accuracy | ± 10C |
| Control System | Imported Microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value. |

| Models | Capacity Liters | Internal size L x W x H cm | Standard Model with lid (Amt. in Rs.) | Inner & Outer SS 304 With lid (Amt. in Rs.) | GMP Model with lid (Amt. in Rs.) | Packing Charge |
| --- | --- | --- | --- | --- | --- | --- |
| TI-611 | 14 | 30 x 30 x 15 | 68,800 | 79,800 | 82,400 | 2,500 |
| TI-612 | 20 | 45 x 30 x 15 | 73,500 | 86,100 | 88,800 | 2,500 |
| TI-613 | 30 | 44 x 38 x 18 | 81,700 | 95,600 | 98,500 | 2,800 |
| TI-614 | 40 | 45 x 30 x 30 | 84,500 | 97,700 | 1,01,400 | 3,500 |
| TI-615 | 50 | 45 x 35 x 30 | 91,800 | 1,08,200 | 1,11,900 | 4,000 |

OPTIONALS :-

| 1 | Safety digital temperature controller with accessories | 6,000 |
| --- | --- | --- |
| 2 | Digital Timer 0-999 ( Hr / Min / Sec ) | 4,000 |
| 3 | Low water Level Alarm : Float switch to cut off the heater supply in case of low water level in the tank and give Audio & Visual alarm. | 4,500 |

Requirement for Installation:

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply. MCB with proper Earthing. (230V + 5%, 50Hz, 1 Phase, 20 Amp). It is

recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures.

CONSTANT TEMPERATURE WATER BATH  With Circulation:-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | Double walled construction with inner S.S. argon welded one piece pot which avoids leakages. |
| Insulation | 2” Thick Glass Wool Insulation ensures temperature stability & reduced energy consumption. |
| Water Circulation | Circulation pump is provided for better uniformity of Temperature instead of stirrer. Side mounted circulating pump 18 ltr. /min capacity, single phase, 230 Volts. |
| Heating System | S.S. Tubular heater (water immersion type) is used at bottom for better heat distribution. |
| Lid | Top mounted removable lid. |
| Features | - Compact table top design. - Circulation pump use for Drain (Do not dry run circulations pump). |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 argon welded one piece pot Outer Galvanized Iron (GI) sheet with powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 argon welded one piece pot Outer S.S. 304. |
| G.M.P. Model | Inside S.S. 316 argon welded one piece pot Outer S.S. 304 |
| Temperature Range | 50 C above ambient to 900 C |
| Temperature Accuracy | ± 0.50 C |
| Control System | Imported microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |

| Models | Capacity Liters | Internal size L x W x H cm | Standard Model with Lid (Amt. in Rs.) | Inner & Outer SS 304 with Lid (Amt. in Rs.) | GMP Model with Lid (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| TI-601 | 14 | 30 x 30 x 15 | 35,300 | 39,300 | 41,200 | 2,000 |
| TI-602 | 20 | 45 x 30 x 15 | 38,300 | 43,400 | 46,200 | 2,300 |
| TI-603 | 30 | 44 x 38 x 18 | 40,400 | 46,200 | 49,600 | 2,800 |
| TI-604 | 40 | 45 x 30 x 30 | 42,700 | 49,000 | 52,500 | 3,300 |
| TI-605 | 50 | 45 x 35 x 30 | 45,600 | 52,500 | 56,700 | 3,300 |

OPTIONALS :-

| 1 | Low water Level Alarm : Float switch to cutoff the heater supply in case of low water level in the tank and give Audio & Visual alarm. | 4,500 |
| --- | --- | --- |

Requirement for Installation:

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures

SHAKER BATHS  :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Shaking model | Horizontal |
| Construction | Double walled construction |
| Insulation | 2” Thick Glass Wool Insulation |
| Heating System | S.S. Tubular heater (water immersion type) is used at bottom for better heat distribution. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 seamless pot (argon welded) which avoid water leakage problems Outer Galvanized Iron (GI) sheet with powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 seamless pot (argon welded) which avoid water leakage problems & Outer S.S. 304. |
| GMP Model | Inside S.S. 316 seamless pot (argon welded) which avoid water leakage problems & Outer S.S. 304. |
| Temperature Range | 50C above ambient to 950C |
| Temperature Accuracy | ± 0.50C |
| Control System | Imported microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |
| Shaking Speed | 40 to 120 RPM |
| Speed indicator | Digital RPM Controller & Indicator. |

| Models | Capacity | Standard Model (Amt. in Rs.) | Inner & Outer SS 304 (Amt. in Rs.) | GMP Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-242 A | 9 flask of 100 ml | 75,300 | 79,600 | 82,400 | 3,000 |
| TI-242 B | 12 flask of 250 ml | 78,800 | 84,700 | 87,900 | 4,000 |

| Optional | Digital Timer 0-999 ( Hr / Min / Sec ) with Alarm | 4,500 |
| --- | --- | --- |

WATER BATH (Electric, Thermostatic)  :

Make :  “TEMPO”

Double walled 50C above ambient to 950C. Accuracy +20C with drop  on  lid  and rings.

(With 4 holes of 7.5 cm dia and 1 hole of 12.5 cm dia.)

| Models | HSN Code | Inner Dimensions (L x W x D Cm) | Rating | Price Rs. | Packing Charges |
| --- | --- | --- | --- | --- | --- |
| TI-240 | 8419 | 30 x 25 x 12.5 | 1.0 KW | 15,600 | 1,000 |

WATER BATH   (Electric, Thermostatic)  :

Make :  “TEMPO”

HSN Code :  8419

Double walled 50C above ambient to 950C. Accuracy +2 C with drop on lid and rings.

STANDARD Model – Inside S.S. 304 pot (argon welded) which avoid water leakage problems.

Outer Galvanized Iron (GI).

Inner & Outer SS 304 - Inside S.S. 304 mirror finish seamless pot (argon welded) which avoid water leakage

problems.  Outer S.S. 304 with dull smooth buff finish

GMP Model -  Inside S.S. 316 mirror finish seamless pot (argon welded) which avoid water leakage problems.

Outer S.S. 304 with dull smooth buff finish.

| Models | Inner Dimensions L x W x D Cm | Rating | Standard Model | Inner & Outer SS 304 | GMP Model | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| TI-241 A | 35 x 25 x 10 With 6 holes of 7.5 Cm dia. | 1.0 KW | 16,200 | 18,900 | 21,000 | 1,200 |
| TI-241 B | 40 x 30 x 10 With 12 holes of 7.5 cm dia. | 1.5 KW | 19,200 | 22,300 | 24,800 | 1,500 |

| Optional | Imported auto tuned PID Controller cum Indicator with PT-100 Sensor | 5,600 |
| --- | --- | --- |

SEROLOGICAL WATER BATH :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | Double walled construction with inner S. S. argon welded one piece pot which avoids leakages. |
| Insulation | 2” Thick Glass Wool Insulation ensures temperature stability & reduced energy consumption. |
| Heating System | S. S. Tubular heater (water immersion type) is used at bottom for better heat distribution. |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 argon welded one piece pot Outer Galvanized Iron (GI) sheet with powder coated. |
| G.M.P. Model | Inside S.S. 316 argon welded one piece pot & Outer S.S. 304. |
| Pyramidal Lid | With Pyramidal Lid |
| Temperature Range | 50 C above ambient to 600 C |
| Temperature Accuracy | ± 0.5 0C |

| Models | Test Tube Racks Capacity | Capacity Liters | Rating Watt | Internal size L x W x H cm | Standard Model (Amt. In Rs.) | G.M.P. Model (Amt. In Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TI-150 A | 1 | 4.2 | 250 | 25 x 13 x 13 | 13,200 | 17,900 | 750 |
| TI-150 B | 2 | 8 | 500 | 25 x 25 x 13 | 16,000 | 24,200 | 1,200 |
| TI-150 C | 3 | 12 | 1000 | 37 x 25 x 13 | 19,600 | 29,400 | 1,400 |

NOTE : THERMOMETER NOT SUPPLIED WITH WATER BATH

TEST TUBE RACKS ARE SUPPLIED AT AN EXTRA COST AS MENTIONED IN

OPTIONS

OPTIONALS :-

| 1 | Imported auto tuned PID Controller cum Indicator with PT-100 Sensor |  | 5,600 |
| --- | --- | --- | --- |
| 2 | Test tube racks : 40 Tubes in one test tube rack. Useful for Size of test tube - Dia – 12 mm | S.S. 316 | 1,200 |
|  |  | S.S. 304 | 1,000 |
| 3 | Low water Level Alarm: Float switch to cutoff the heater supply in case of low water level in the tank and give Audio & Visual alarm. |  | 4,500 |

Requirement for Installation:

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V 5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures

MBRT WATER BATH :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | Double walled construction with inner S. S. argon welded one piece pot which avoids leakages. |
| Insulation | 2” Thick Glass Wool Insulation ensures temperature stability & reduced energy consumption. |
| Heating System | S. S. Tubular heater (water immersion type) is used at bottom for better heat distribution. |
| Stirrer | Electrically operated stirrer with S.S. blade |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 argon welded one piece pot Outer Galvanized Iron (GI) sheet with powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 argon welded one piece pot. Outer S.S. 304 |
| G.M.P. Model | Inside S.S. 316 argon welded one piece pot. Outer S.S. 304. |
| Flat Lid | With Flat Lid |
| Temperature Range | 50 C above ambient to 980 C |
| Temperature Accuracy | ± 0.5 0C |
| PID Controller | Imported auto tuned PID Controller cum Indicator with PT-100 Sensor |

| Models | Test Tube Racks Cap. | Cap. Liters | Internal size L x W x H in cm | Standard Model (Amt. In Rs.) | Inner & Outer 304 (Amt. In Rs.) | G.M.P.Model (Amt. In Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- | --- |
| TI-160 A | 2 | 14 | 40 x 20 x 18 | 35,200 | 38,800 | 41,300 | 2,000 |
| TI-160 B | 4 | 25 | 40 x 35 x 18 | 37,300 | 42,300 | 45,200 | 3,000 |
| TI-160 C | 6 | 35 | 55 x 35 x 18 | 39,400 | 45,200 | 48,300 | 3,500 |

NOTE : 1) Above price are excluding test tube rack cost.

2) Thermometer not supplied with water bath.

OPTIONALS :-

| 1 | Pyramidal Lid | 3,000 |
| --- | --- | --- |
| 2 | Test tube racks : 30 Tubes in one test tube rack. Useful for Size of test tube – 15 x 150 mm OR 18 x 150 mm | 1,500 |
| 3 | Low water Level Alarm: Float switch to cutoff the heater supply in case of low water level in the tank and give Audio & Visual alarm. | 4,500 |
| 4 | Glass window on one side | 10,000 |

SPARE PARTS (Keep in Stock)

| 1 | Imported auto tuned PID Controller cum Indicator with PT-100 Sensor | 5,600 |
| --- | --- | --- |
| 2 | Electrically operated stirrer with S.S. blade | 9,000 |

Requirement for Installation:

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V 5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage stabilizer(KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures

THAWING WATER BATH :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Certification | CE Certification |
| Construction | Double walled construction with inner S.S. argon welded one piece pot which avoids leakages. |
| Insulation | 2” Thick Glass Wool Insulation ensures temperature stability & reduced energy consumption. |
| Water Circulation | Circulation pump is provided for better uniformity of Temperature instead of stirrer. Side mounted circulating pump 18 ltr. /min capacity, single phase, 230 Volts. |
| Heating System | S.S. Tubular heater (water immersion type) is used at bottom for better heat distribution. |
| Lid | Top mounted removable lid. |
| Features | - Compact table top design. - Circulation pump use for Drain (Do not dry run circulations pump). |
| Supply | 230V AC, Single Phase, 50Hz |

| SPECIFICATIONS : |  |
| --- | --- |
| Standard Model | Inside S.S. 304 argon welded one piece pot Outer Galvanized Iron (GI) sheet with powder coated. |
| Inner & Outer SS 304 | Inside S.S. 304 argon welded one piece pot & Outer S.S. 304 |
| G.M.P. Model | Inside S.S. 316 argon welded one piece pot & Outer S.S. 304 |
| Temperature Range | 50 C above ambient to 900 C |
| Temperature Accuracy | ± 0.5 0C |
| Control System | Imported microprocessor based auto tuned PID controller with CE mark & dual display of set value & process value for precise control of temperature. |

| Models | Capacity Liters | Internal size L x W x H cm | Standard Model with Lid (Amt. in Rs.) | Inner & Outer SS 304 with Lid (Amt. in Rs.) | GMP Model with Lid (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| TI-631 | 14 | 30 x 30 x 15 | 35,300 | 39,300 | 41,200 | 2,000 |
| TI-632 | 20 | 45 x 30 x 15 | 38,300 | 43,400 | 46,200 | 2,300 |
| TI-633 | 30 | 44 x 38 x 18 | 40,400 | 46,600 | 49,600 | 2,800 |

OPTIONALS :-

| 1 | Low water Level Alarm : Float switch to cut off the heater supply in case of low water level in the tank and give Audio & Visual alarm. | 4,500 |
| --- | --- | --- |

Requirement for Installation:

- Place the chamber where ventilation is good.

- The chamber must be operated in an ambient of 220C ~ 300C.

- Leave minimum 1ft.space from the sides and leave 2ft. space from the backside.

- Place the chamber horizontally leveled.

- Give stabilized electric supply (230V +5%, 50Hz, 1 Phase, 20 Amp). It is recommended to use a voltage stabilizer (KVA).

- Proper grounding of the chamber to get satisfactory operation of chamber.

- Do not test any product in the chamber, which is capable of generating combustible mixtures

STEAM POT   :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8419 |
| Application | Ideal for Digestion media in Conical flaks (Upto 2000 ml) As well as for Inspissation. |
| Construction | - Fully stainless steel double walled contrition with insulation of preventing energy losses. - Hinges lid with assisted lift for easy opening as well as effortless loading & removal of flaks. - SS wire mesh heater cover for efficient steam movement and stable placement of media flaks. - Drain valve for draining chamber water. - Open port at top for venting of steam. - Validation port provided as part of standard supply. |
| Heater | Special energy efficient tubular heaters |

| SPECIFICATIONS : |  |
| --- | --- |
| Inner & Outer SS 304 | Inside S.S. 304 Outer S.S. 304 |
| Features | - Water level indicating device - Safety thermostatic cut-off (self-resetting type) for assured low water detection – with visual indication. |
| Temperature controller | Digital Temperature controller with Timer |
| Temperature Range | 50C above ambient to 1000 C |
| Accuracy | + 10C |
| Resolution | 0.10C |

| Models | Capacity Liters | Internal Tank size L x B x H cm | Amount in Rs. (EACH) | Packing Charges |
| --- | --- | --- | --- | --- |
| TI – 672 A | 24 | 40 x 30 x 20 | 65,100 | 3,500 |
| TI - 672 B | 84 | 60 x 40 x 35 | 94,500 | 5,000 |

OPTIONALS :-

| Sr. No. | Description | Amount in Rs. | Packing Charges |
| --- | --- | --- | --- |
| 1 | S.S. Stand with 2 feet height for TI – 672 A | 16,000 | 2,000 |
| 2 | S.S. Stand with 2 feet height for TI – 672 B | 19,000 | 3,000 |

SOXHLET EXTRACTION HEATERS :-

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Certification | CE Certification |
| Usage | Are used in laboratories for Solvent Extraction under reflux. Typical applications are determination of residual oil content in oil cake, assay of crude herbal drug sources etc. The apparatus is suitable to heat Soxhlet flasks of 100-350 ml capacity. |
| Heating System | The heating element is attached to knitted fabric to provide flexible support to the bottom of the flask. Glass wool insulation fills the gap between the metal casing and the heater. |
| Features | In event of spillage, the solvent will either evaporate or escape through the draining hole below each element. The housing has two retort rods and a cross bar on which the extraction apparatus is fixed. The Apparatus has built in energy regulator which controls all the heaters. The heating elements have individual ON/OFF switch and pilot lamp to indicate working. |
| Supply | 230V AC, Single Phase |

| Models | Heaters | Soxhlet Flasks Capacity | Rating | Temperature | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- | --- | --- |
| TI-184 IRB | 4 | 100 - 350 ml | 500 W | 4000C | 9,800 | 850 |
| TI-184 IRC | 6 | 100 - 350 ml | 750 W | 4000C | 13,000 | 1,100 |

FLEXIBLE INSULATED HEATING CORD (400°C) (HIGH WATTAGE)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | - Designed for use on small tubes, vessels, or any application where space is limited. - Heating cords are used when heating tapes are too bulky or cumbersome |
| Features | - Ideal for laboratory, R&D, gas tubing, production, maintenance, temporary heat, and more... - These 6 mm diameter cords solve a wide range of limited-space heating applications in laboratory, production, Industrial applications. - Rapid thermal response. |
| Temperature | 400°C High Temperature Capability for continuous use |

| Type No | Length (mm) | Wattage | Voltage | Amt. in Rs. |
| --- | --- | --- | --- | --- |
| TI-641-INS-H1 | 1000 mm | 200 W | 230 V | 650 |
| TI-641-INS-H2 | 2000 mm | 400 W | 230 V | 860 |
| TI-641-INS-H3 | 3000 mm | 600 W | 230 V | 1,150 |
| TI-641-INS-H4 | 4000 mm | 800 W | 230 V | 1,560 |
| TI-641-INS-H5 | 5000 mm | 1000 W | 230 V | 2,200 |

Warranty – Heating cords are not supported in warranty.

OPTIONALS :-

| Cat. No.: FLX - DTC | Descriptions | Amt. in Rs. |
| --- | --- | --- |
| HSN Code 9032 | Digital Temp. Controller: Temperature Controller System (Digital)can be provided for display, monitoring and safe operation. | 7,200 |

HEATING CORDS :- (LOW WATTAGE)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | - Designed for use on small tubes, vessels, or any application where space is limited. - Heating cords are used when heating tapes are too bulky or cumbersome |
| Features | - Ideal for laboratory, R&D, gas tubing, production, maintenance, temporary heat, and more... - These 6 mm diameter cords solve a wide range of limited-space heating applications in laboratory, production, Industrial applications. |
| Temperature | 400°C High Temperature Capability for continuous use |

| Models | Length (mm) | Wattage | Voltage | Amt. in Rs. |
| --- | --- | --- | --- | --- |
| TI-64 A | 2500 x 3 Cords | 125 x 3 | 230 V | 2,500 |
| TI-64 B | 5000 x 3 Cords | 250 x 3 | 230 V | 3,900 |
| TI-64 C | 10000 x 3 Cords | 500 x 3 | 230 V | 6,100 |
| TI-64 D | 20000 x 3 Cords | 1000 x 3 | 230 V | 8,500 |

Warranty – Heating cords are not supported in warranty.

NOTE : FOR DOUBLE COATING WILL BE RS. 1,000/- EXTRA.

OPTIONALS :-

| Cat. No.: FLX - DTC | Descriptions | Amt. in Rs. |
| --- | --- | --- |
| HSN Code 9032 | Digital Temp. Controller: Temperature Controller System (Digital)can be provided for display, monitoring and safe operation. | 7,200 |

SILICON HEATING CORD (100°C)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | Heating Cords with Silicone Rubber Cover provided for use on conductive and moisture surface. |
| Features | - Safe for use with metal and electrically conductive surfaces. - Widely used in deep freezer ,refrigerators etc |
| Temperature | Maximum exposure temperature - (Removable & reusable use 100°C) |

| Models | Size in Metres | Wattage | Voltage | Amt. in Rs. |
| --- | --- | --- | --- | --- |
| TI-641-SIL-L1 | 1000 mm | 40 W | 230 V | 900 |
| TI-641-SIL-L2 | 2000 mm | 80 W | 230 V | 1,650 |
| TI-641-SIL-L3 | 3000 mm | 120 W | 230 V | 2,300 |
| TI-641-SIL-L4 | 4000 mm | 160 W | 230 V | 3,050 |
| TI-641-SIL-L5 | 5000 mm | 200 W | 230 V | 3,800 |

Warranty – Heating cords are not supported in warranty.

OPTIONALS :-

| Cat. No.: FLX - DTC | Descriptions | Amt. in Rs. |
| --- | --- | --- |
| HSN Code 9032 | Digital Temp. Controller: Temperature Controller System (Digital)can be provided for display, monitoring and safe operation. | 7,200 |

HEATING MANTLES  :-

| Make | “Tempo” |
| --- | --- |
| HSN Code | 8514 |
| Construction | Standard Model: Body is made of Aluminium pot with powder coated which avoid rusting. |
| Insulation | High quality glasswool insulation |
| Maximum Temp. | Maximum surface temperature 350 0C. |
| Temp. Controller | Energy Regulator and Pilot Lamp. |
| Heating Systems | The heating cords are stitched in glass yarn knitted cloth. The glass yarn acts as good electrical insulator, and dissipates the heat very effectively. Booster heaters – in bigger sizes |
| Usage | Heating Mantles are mainly used for heating/boiling solutions in round bottom flasks |
| Volt | 230 Volt |

| Models | Flask Size | Rating | Standard Model (Amt. in Rs.) | Packing Charges |
| --- | --- | --- | --- | --- |
| TI-66 RA | 50 - 250 ml | 120 W | 4,200 | 500 |
| TI-66 RB | 100 - 500 ml | 160 W | 5,100 | 500 |
| TI-66 RC | 250 - 1000 ml | 300 W | 5,800 | 500 |
| TI-66 RD | 500 - 3000 ml | 2 x 250 W | 7,200 | 600 |
| TI-66 RE | 1L - 5 Ltr. | 2 x 300 W | 9,000 | 700 |

FLEXIBLE INSULATED HEATING CORD (250°C)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | - Designed for use on small tubes, vessels, or any application where space is limited. - Heating cords are used when heating tapes are too bulky or cumbersome |
| Features | - Ideal for laboratory, R&D, gas tubing, production, maintenance, temporary heat, and more... - These 6 mm diameter cords solve a wide range of limited-space heating applications in laboratory, production, Industrial applications. - Rapid thermal response. |
| Temperature | 400°C High Temperature Capability for continuous use |

| Type No | Length (mm) | Wattage | Voltage | Amt. in Rs. |
| --- | --- | --- | --- | --- |
| TI-641-INS-H1 | 1000 mm | 200 W | 230 V | 600 |
| TI-641-INS-H2 | 2000 mm | 400 W | 230 V | 900 |
| TI-641-INS-H3 | 3000 mm | 600 W | 230 V | 1,100 |
| TI-641-INS-H4 | 4000 mm | 800 W | 230 V | 1,500 |
| TI-641-INS-H5 | 5000 mm | 1000 W | 230 V | 2,100 |

OPTIONALS :-

| Cat. No.: FLX - DTC | Descriptions | Amt. in Rs. |
| --- | --- | --- |
| HSN Code 9032 | Digital Temp. Controller: Temperature Controller System (Digital)can be provided for display, monitoring and safe operation. | 7,600 |

HEATING CORDS :- (LOW WATTAGE) (400°C)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | - Designed for use on small tubes, vessels, or any application where space is limited. - Heating cords are used when heating tapes are too bulky or cumbersome |
| Features | - Ideal for laboratory, R&D, gas tubing, production, maintenance, temporary heat, and more... - These 6 mm diameter cords solve a wide range of limited-space heating applications in laboratory, production, Industrial applications. |
| Temperature | 400°C High Temperature Capability for continuous use |

| Models | Length (mm) | Wattage | Voltage | Amt. in Rs. |
| --- | --- | --- | --- | --- |
| TI-64 A | 2500 x 3 Cords | 125 x 3 | 230 V | 2,700 |
| TI-64 B | 5000 x 3 Cords | 250 x 3 | 230 V | 4,100 |
| TI-64 C | 10000 x 3 Cords | 500 x 3 | 230 V | 6,500 |
| TI-64 D | 20000 x 3 Cords | 1000 x 3 | 230 V | 9,000 |

OPTIONALS :-

| Cat. No.: FLX - DTC | Descriptions | Amt. in Rs. |
| --- | --- | --- |
| HSN Code 9032 | Digital Temp. Controller: Temperature Controller System (Digital)can be provided for display, monitoring and safe operation. | 7,600 |

NOTE : FOR DOUBLE COATING WILL BE RS. 1,000/- EXTRA.

HEAVY INSULATED HEATING TAPES (400°C) (HIGH WATT DENSITY)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | - Lead Wires on one end - Spiral Heating Element design to give better temperature uniformity |
| Insulation | Heavy Braided Fiber Glass Outer Cover |
| Features | - Robust & Flexible Element - Low Watt Density, enhancing longer life through better distribution of heat. (Lower watts per cm of wire) - 25 mm wide tapes may be easily applied to diameters as small as 12mm |
| Temperature | 350OC High Temperature Capability for continuous use |

| Models | Size (Width x Length mm) | Wattage | Voltage | Amt. in Rs. |
| --- | --- | --- | --- | --- |
| TI-681-HIT-H2 | 25 mm x 600 mm | 310 W | 230 V | 2,200 |
| TI-681-HIT-H4 | 25 mm x 1200 mm | 620 W | 230 V | 3,900 |
| TI-681-HIT-H6 | 25 mm x 1800 mm | 930 W | 230 V | 5,700 |
| TI-681-HIT-H8 | 25 mm x2400 mm | 1240 W | 230 V | 7,800 |

HEAVY INSULATED HEATING TAPES 350°C (MEDIUM WATT DENSITY)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | - Lead Wires on one end - Spiral Heating Element design to give better temperature uniformity |
| Insulation | Heavy Braided Fiber Glass Outer Cover |
| Features | - Robust & Flexible Element - Low Watt Density, enhancing longer life through better distribution of heat. (Lower watts per cm of wire) - 25 mm wide tapes may be easily applied to diameters as small as 12mm |
| Temperature | 350OC High Temperature Capability for continuous use |

| Cat. No. | Size (Width x Length mm) | Wattage | Voltage | Amt. in Rs. |
| --- | --- | --- | --- | --- |
| TI-681-HIT-M2 | 25 mm x 600 mm | 210 W | 230 V | 1,900 |
| TI-681-HIT-M4 | 25 mm x 1200 mm | 420 W | 230 V | 3,300 |
| TI-681-HIT-M6 | 25 mm x 1800 mm | 630 W | 230 V | 4,700 |
| TI-682-HIT-M8 | 25 mm x 2400 mm | 832 W | 230 V | 6,500 |
| TI-682-HIT-M10 | 25 mm x 3000 mm | 1040 W | 230 V | 10,000 |
| TI-681-HIT-M16 | 25 mm x 5 Meter | 1670 W | 230 V | 10,300 |
| TI-681-HIT-M20 | 25 mm x 6 Meter | 1800 W | 230 V | 10,800 |
| TI-681-HIT-M26 | 25 mm x 8 Meter | 2000 W | 230 V | 11,500 |
| TI-681-HIT-M32 | 25 mm x 10 Meter | 2200 W | 230 V | 12,600 |
| TI-681-HIT-M38 | 25 mm x 12 Meter | 2400 W | 230 V | 13,200 |

HEAVY INSULATED HEATING TAPES 200°C (LOW WATT DENSITY)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | - Lead Wires on one end - Spiral Heating Element design to give better temperature uniformity |
| Insulation | Heavy Braided Fiber Glass Outer Cover |
| Features | - Robust & Flexible Element - Low Watt Density, enhancing longer life through better distribution of heat. (Lower watts per cm of wire) - 25 mm wide tapes may be easily applied to diameters as small as 12mm |
| Temperature | 350OC High Temperature Capability for continuous use |

| Models | Size (Width x Length mm) | Wattage | Voltage | Amt. in Rs. |
| --- | --- | --- | --- | --- |
| TI-681-HIT-L2 | 25 mm x 600 mm | 80 W | 230 V | 1,600 |
| TI-681-HIT-L4 | 25 mm x 1200 mm | 160 W | 230 V | 2,800 |
| TI-681-HIT-L6 | 25 mm x 1800 mm | 230 W | 230 V | 4,000 |
| TI-681-HIT-L8 | 25 mm x 2400 mm | 310 W | 230 V | 5,500 |
| TI-681-HIT-L10 | 25 mm x 3000 mm | 385 W | 230 V | 8,300 |
| TI-681-HIT-L12 | 25 mm x 3200mm | 468 W | 230 V | 10,500 |
| TI-681-HIT-L16 | 25 mm x 4800mm | 624 W | 230 V | 15,100 |
| TI-681-HIT-L32 | 25 mm x 10 Meter | 1240 W | 230 V | 13,000 |

STANDARD INSULATED HEATING TAPES 350°C (Without Sleeve Tape)

(MEDIUM WATT DENSITY)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | - Lead Wires on one end - Spiral Heating Element design to give better temperature uniformity. - Robust & Flexible Element. |
| Features | - Lower Watt Density, enhancing longer life through better distribution of heat. (Lower watts per cm of wire) - 25 mm wide tapes may be easily applied to diameters as small as 12 mm. |
| Temperature | 350°C High Temperature Capability for continuous use |

| Models | Size (Wdth x Length mm) | Wattage | Voltage | Amt. in Rs. |
| --- | --- | --- | --- | --- |
| TI-681-SIT-L2 | 25 mm x 600 mm | 210 W | 230 V | 1,300 |
| TI-681-SIT-L4 | 25 mm x 1200 mm | 420 W | 230 V | 2,200 |
| TI-681-SIT-L6 | 25 mm x 1800 mm | 630 W | 230 V | 3,200 |
| TI-682-SIT-M2 | 25 mm x 600 mm | 210 W | 230 V | 1,500 |
| TI-682-SIT-M4 | 25 mm x 1200 mm | 420 W | 230 V | 2,600 |
| TI-682-SIT-M6 | 25 mm x 1800 mm | 630 W | 230 V | 3,800 |

OPTIONALS :-

| Cat. No.: FLX - DTC | Descriptions | Amt. in Rs. |
| --- | --- | --- |
| HSN Code 9032 | Digital Temp. Controller: Temperature Controller System (Digital)can be provided for display, monitoring and safe operation. | 7,600 |

SILICON HEATING TAPE (180°C)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | - Lead Wires on one end - Extruded Silicone Rubber Outer Cover - Moisture Resistant Construction - The Silicone Rubber Extrusion uses a low durometer with heat additives to produce the maximum exposure and continuous operating temp. capability, Longer Life, Lower Watt Density. |
| Features | - Lower watts per inch of wire, enhances longer. Life through better distribution of heat. - Ultimate flexibility. 25mm wide tapes may, be easily applied to diameters as small as12mm. - Safe to use on conductive surfaces such as metal tubing. |
| Temperature | Max operation Temperature 200°C |

| Models | Size (Wdth x Length mm) | Wattage | Voltage | Amt. in Rs. |
| --- | --- | --- | --- | --- |
| TI-681-SIL-L2 | 25 mm x 600 mm | 100 W | 230 V | 3,100 |
| TI-681-SIL-L4 | 25 mm x 1200 mm | 200 W | 230 V | 5,300 |
| TI-681-SIL-L6 | 25 mm x 1800 mm | 310 W | 230 V | 8,300 |

OPTIONALS :-

| Cat. No.: FLX - DTC | Descriptions | Amt. in Rs. |
| --- | --- | --- |
| HSN Code 9032 | Digital Temp. Controller: Temperature Controller System (Digital)can be provided for display, monitoring and safe operation. | 7,600 |

SILICA  TAPE (600°C)

| Make | “TEMPO” |
| --- | --- |
| HSN Code | 8514 |
| Construction | - Lead Wires on one end - Extruded Silicone Rubber Outer Cover - Moisture Resistant Construction - The Silicone Rubber Extrusion uses a low durometer with heat additives to produce the maximum exposure and continuous operating temp. capability, Longer Life, Lower Watt Density. |
| Features | - Lower watts per inch of wire, enhances longer. Life through better distribution of heat. - Ultimate flexibility. 25mm wide tapes may, be easily applied to diameters as small as12mm. - Safe to use on conductive surfaces such as metal tubing. |
| Temperature | Max operation Temperature 200°C |

| Models | Size (Wdth x Length mm) | Wattage | Voltage | Amt. in Rs. |
| --- | --- | --- | --- | --- |
| TI-681-HST-M2 | 25 mm x 600 mm | 450 W | 230 V | 4,200 |
| TI-681-HST-M4 | 25 mm x 1200 mm | 900 W | 230 V | 7,600 |
| TI-681-HST-M6 | 25 mm x 1800 mm | 1350 W | 230 V | 11,000 |

OPTIONALS :-

| Cat. No.: FLX - DTC | Descriptions | Amt. in Rs. |
| --- | --- | --- |
| HSN Code 9032 | Digital Temp. Controller: Temperature Controller System (Digital)can be provided for display, monitoring and safe operation. | 7,600 |